## Sheet1
Export Reception file
LAARC V2.0, August 2020
[138 empty] | 1 | [10 empty] | 2 | [5 empty] | 3 | [6 empty] | 4 |  |  |  | 5 | [5 empty] | 6 | [7 empty] | 7 |  |  |  |  | 8 |  |  |  |  | 9 |  |  |  | 10 | [14 empty] | 11 | [5 empty] | 12 | [9 empty] | 13 |  |  | 14
| General information | [17 empty] | Manual AST in use |  |  |  |  | Autom. AST in use |  |  |  | Organisms/Chromagar | [5 empty] | Genes/PCR | [5 empty] | QMS Mentoring pgm | [5 empty] | Accreditation/Certification | [6 empty] | TEST MENU AND ANNUAL CULTURE WORKLOAD | [20 empty] | ANNUAL AST WORKLOAD | [5 empty] | Biochemical identification QC score | [24 empty] | Biochemical identification SOP score | [24 empty] | STAFFING | [12 empty] | General equipment | [20 empty] | Media Prep equipment | [5 empty] | Autom. Instruments |  |  |  |  | FLAG # |  |  | GI | Section Scores/Performance Heatmap | [13 empty] | laboratory facility | [10 empty] | LIS | [5 empty] | Data management | [6 empty] | Quality assurance |  |  |  | Media QC | [5 empty] | ID QC | [7 empty] | AST QC |  |  |  |  | Specimen collection |  |  |  |  | Processing |  |  |  | Identification | [14 empty] | Basic AST | [5 empty] | Expert rules | [9 empty] | AST panels/policy |  |  | Biosafety |  |  |  | Red Flag | [100 empty] | Training opportunities | [9 empty] | System flag | [23 empty] | 1- Facility questions | [20 empty] | 2- LIS questions | [44 empty] | 5- Media questions | [12 empty] | 6-ID_Q |  | 9-Proc-Q | 10-Identification questions | [34 empty] | 11- Basic AST questions | [12 empty] | 12-AST expert rules questions | [32 empty] | 13-ASTP Q
Full name | Laboratory/Hospital name | Address | City | Province | District | Country | Assessor 1 (last name and affiliation) | Assessor 2 (last name and affiliation) | Assessor 3 (last name and affiliation) | Date of assessment (dd/mm/yyyy) | Altitude | Latitude | Longitude | Primary Laboratory funding sources | Primary Laboratory affiliation | Laboratory/Facility Level | Service level of the Facility | Number of beds of the Facility | Disk diffusion | Gradient Strip (e.g., Etest/Liofilchem) | Broth microdilution (96-well tray) | Broth macrodilultion (tube method) | Agar dilution | Vitek | Phoenix | Microscan | Other, please specify in comments. | ESBL producers | CRE/Carbapenemases | MRSA | VRE | Colistin resistance | Other, please specify in comments. | ESBLs | Carbapenemases | mecA | vanA/vanB | mcr-1 | Other, please specify in comments. | Lab enrolled in the SLIPTA program? | Most recent certification awarded? | what star level of the latest SLIPTA audit? | Lab enrolled in the WHO LQSI program? | Overall % score for the 4 phases? | Other mentoring program? | Blood cultures | Stool cultures | Urine cultures | Organism Identification | Antibiotic Susceptibility Testing | Other such as Gram staining? | Accreditation/certification body | Blood Cultures | Urine Cultures | Stool Cultures | Salmonella and/or Shigella | Vibrio cholerae | Yersinia enterocolitic | Campylobacter jejuni | Enterohemorrhagic/Enterotoxic E.coli (e.g., O157:H7) | Respiratory Cultures (not TB) | Wound Cultures | Cerebrospinal Fluid Cultures | Sterile Body Fluid Cultures | Genital Cultures | Anaerobic Cultures | Fungal Cultures (Yeast) | Fungal Cultures (Mold) | MRSA screen (nares, axilla, groin) | VRE screen (rectal swab) | CRE screen (rectal swab) | ID/AST of isolates referred from other labs | Other cultures of local imoprtance | Automated AST instrument | Disk diffusion | Gradient Strip (e.g., Etest/Liofilchem) | Broth microdilution (96-well tray) | Broth macrodilultion (tube method) | Agar dilution | Catalase (H2O2) | Coagulase plasma | Staph latex agglutination | Staph chromagar | DNase | PYR | Optochin ("P") disk | Bile solubility (deoxycholate) | Strep. pneumo latex agglutination | Oxidase | Indole reagents | Methyl Red | Voges-Proskauer | Citrate | Triple Sugar Iron agar or Kligler's Iron Agar | Urease | Motility | Lysine Iron Agar (LIA) or Lysine decarboxylase (LDC) | Glucose or Dextrose Oxidative-Fermentative (OF) test | Nitrate reduction | Gelatin hydrolysis | Chloramphenicol resistance (disk) | Growth at 42°C | Shigella serology | Salmonella serology | Catalase (H2O2) | Coagulase plasma | Staph latex agglutination | Staph chromagar | DNase | PYR | Optochin ("P") disk | Bile solubility (deoxycholate) | Strep. pneumo latex agglutination | Oxidase | Indole reagents | Methyl Red | Voges-Proskauer | Citrate | Triple Sugar Iron agar or Kligler's Iron Agar | Urease | Motility | Lysine Iron Agar (LIA) or Lysine decarboxylase (LDC) | Glucose or Dextrose Oxidative-Fermentative (OF) test | Nitrate reduction | Gelatin hydrolysis | Chloramphenicol resistance (disk) | Growth at 42°C | Shigella serology | Salmonella serology | Advanced degree in Med. Microbiology/Med. Lab. (PhD, MD) | Advanced degree, other concentration (PhD, MD) | PG Master's degree in Microbiology or Med. Lab. Sciences | Postgraduate Master's degree, other concentration | Grad. Bachelor's degree in Microbiology or Med. Lab. Sciences | Graduate Bachelor's degree, other concentration | UG Certificate/Diploma in Microbiology or Med. Lab. Sciences | Undergraduate Certificate or Diploma, other concentration | High school/Secondary school diploma | On-the-job training only | Other (specify in comments) | Proportion of staff with Med Micro or Med Lab Training | Proportion of staff with Graduate Bachelor's Degree or more | McFarland QC standards of known densities, including 0.5, not expired | Ruler or caliper with millimeter markings | Bunsen burners or micro-incinerators | Calibrated 1uL or 10uL loops (for plating urine cultures) | Optical densitometer/turbidimeter (for determining McFarland density) | Microliter pipettes (e.g., Eppendorf) | Centrifuge (not used for TB cultures) | Microscope | Thermometers | CO2 incubators | Candle jars | Ambient (non-CO2) incubator | Refrigerator (2-8°C) | Non-defrosting freezer, -20°C | Non-defrosting freezer, -60°C | Non-defrosting freezer, -80°C | Rechargeable desiccants (for storage of open antibiotic disks and strips) | Hot air oven (for drying saturated desiccants) | Biological Safety Cabinet Class IIA | Autoclave for media preparation ("clean" autoclave) | Autoclave for sterlizing waste ("dirty" autoclave) | pH meter | Weighing balance | Conductivity meter | Distillator/reverse osmosis equipment | Hot plate with magnetic stir bar (for mixing powdered media) | Water bath | Blood culture instrument | Instrument for bacterial ID and AST (e.g. Vitek, Phoenix, Microscan) | Instrument for reading disk diffusion (e.g. SIRSCAN, BIOMIC V3, ADAGIO, etc.) | MALDI instrument for organism ID | PCR instrument for antibiotic resistance gene detection | Red Flags | Training Opportunities | System Flags | GENERAL INDICATOR | FACILITY | LAB INFORMATION SYSTEM (Excluded from overall score) | DATA MANAGEMENT | QUALITY ASSURANCE | MEDIA PREPARATION AND QC | ID QC | AST QC | SPECIMEN COLLECTION, TRANSPORT & MANAGEMENT | PROCESSING | IDENTIFICATION | BASIC AST | AST EXPERT RULES | AST PANELS, POLICY AND ANALYSIS | SAFETY | LABORATORY FACILITY | GENERAL EQUIPMENT AVAILABILITY | MEDIA PREPARATION EQUIPMENT AVAILABILITY | EQUIPMENT CALIBRATION RECORDS | THERMOMETERS | TEMPERATURE AND ATMOSPHERE MONITORING | Temperature recorded | Ranges defined | AUTOCLAVE MANAGEMENT | AUTOMATED EQUIPMENT AVAILABILITY AND MAINTENANCE | INVENTORY & STOCK OUTS | DEMOGRAPHIC DATA FIELDS | SPECIMEN DATA FIELDS | WORK CARD DATA FIELDS | AST DATA FIELDS | REPORTS AND DATA TRANSFER CAPABILITIES | INTERFACE CONNECTIVITY | PATIENT AND SPECIMEN IDENTIFICATION | SPECIMEN REQUISITION FORM | ORDER ENTRY | CULTURE OBSERVATIONS | AST RESULTS REPORTING | DATA BACKUP & SECURITY | AMR DATA SHARING | QUALITY STRUCTURE/BASICS | LABORATORY STAFF EDUCATION/TRAINING/COMPETENCY | TOUBLESHOOTING, PROBLEM SOLVING, AND ROOT CAUSE ANALYSES | EXTERNAL QUALITY ASSESSMENT (EQA) | MEDIA PREPARATION SOPs | GENERAL MEDIA PREPARATION | DISTILLED/DEIONIZED WATER PREPARATION | ROUTINE MEDIA QC | MULLER HINTON MEDIA PREPARATION AND QC | BLOOD CULTURE BOTTLES PREPARATION AND QC | GRAM STAIN QC and REAGENT LABELING AND STORAGE | QC OF INDIVIDUAL BIOCHEMICAL METHODS | Positive controls are in use | Negative controls are in use | QC is performed on each new batch/lot number | QC is performed using ATCC or ATCC-derivative strains | QC OF ENTERIC SEROLOGY | QC OF COMMERCIAL ID KITS and AUTOMATED ID SYSTEMS | ROUTINE AST REFERENCE STRAINS | SPECIAL AST REFERENCE STRAINS | QC OF DISC DIFFUSION AST METHODS | QC OF GRADIENT STRIP AST METHODS | QC OF AUTOMATED AST SYSTEMS | SPECIMEN MANAGEMENT | SPECIMEN REJECTION | BLOOD SPECIMEN COLLECTION and TRANSPORT | URINE SPECIMEN COLLECTION and TRANSPORT | STOOL SPECIMEN COLLECTION and TRANSPORT | BLOOD CULTURE PROCESSING | MANUAL BLOOD CULTURE SYSTEMS | URINE CULTURE | STOOL CULTURE for Salmonella and Shigella | CONVENTIONAL ID METHODS - SOP SCORE SUMMARY | Fully implemented*, up-to-date SOP | SOP is readily available** to bench staff | SOP defines QC organisms, frequency, and expected results | SOP provides stepwise instructions for test performance | SOP provides stepwise instructions for test interpretation | STAPHYLOCOCCUS AUREUS, KEY ID METHODS | STAPHYLOCOCCUS AUREUS, OTHER ID METHODS | STREPTOCOCCUS PNEUMONIAE, CONVENTIONAL ID METHODS | ENTEROBACTERIACEAE, CONVENTIONAL ID METHODS | SHIGELLA/SALMONELLA SEROLOGY | ACINETOBACTER SPP, CONVENTIONAL ID METHODS | KIT-BASED ID METHODS | AUTOMATED ID METHODS | IDENTIFICATION FLOWCHARTS | ANTIBIOTIC DISK AND GRADIENT STRIPS MAINTENANCE | INOCULUM PREPARATION | INOCULATION/INCUBATION | READING AST RESULTS | INTERPRETING RESULTS | BREAKPOINTS STANDARDS | EXPERT RULES FOR SALMONELLA | GRAM NEGATIVES & BETA-LACTAM BREAKPOINTS | PHENOTYPIC ESBL TESTING | PHENOTYPIC CARBAPENEMASE TESTING | COLISTIN TESTING | EXPERT RULES FOR STAPH AUREUS | GENERAL CONSIDERATIONS FOR STREP PNEUMONIAE | EXPERT RULES FOR STREP PNEUMONIAE | INDUCIBLE CLINDAMYCIN RESISTANCE TESTING | EXPERT RULES FOR CEREBROSPINAL FLUID | AST PANELS | CUMULATIVE ANTIBIOGRAMS | AST PANEL POLICY | BIOSAFETY EQUIPMENT | BIOSAFETY BEHAVIORS | PERSONAL PROTECTIVE EQUIPMENT | BIOSAFETY DOCUMENTATION AND TRAINING | Indicate whether the lab has the following FUNCTIONAL pieces of equipment. | Biological Safety Cabinet Class IIA | Biological Safety Cabinet Class IIA | If the LIS software automatically interprets zone sizes or MICs, are the breakpoints updated annually? | If the LIS software automatically interprets zone sizes or MICs, are the breakpoints up to date today? | Does the laboratory use the same patient ID numbers assigned by the hospital and/or clinics? | Does the laboratory assign a unique specimen ID number to each specimen received in the lab? | Are specimen numbers assigned in such a way that no two specimens are given the same number during one year? | Are patient results reported if QC of media, ID method, or AST method was not performed? | Are patient results reported if QC of media, ID method, or AST method failed to produce acceptable results? | Is antibiotic disk QC performed before placing newly received lot numbers/shipments into use? (Review QC records to confirm) | Not including new lot QC, how often is antibiotic disk QC performed? (Confirm by reviewing QC records; go back several months) | Is gradient strip QC performed before placing new lot numbers/shipments into use? (Review QC records to confirm) | Not including new lot QC, how often is antibiotic strip QC performed? (Confirm by reviewing QC records; go back several months) | Is QC of the antibiotic cards/trays performed before placing new lot numbers/shipments into use? (Review QC records to confirm) | Not including new lot QC, how often is antibiotic card/tray QC performed? (Confirm by reviewing QC records; go back several months) | Does the lab enforce a two-identifier system? (e.g., both patient name and a numeric identifier must be present on the requisition and on the specimen). | Are unlabeled specimens rejected? | Are mislabeled specimens rejected? | When specimens are rejected, does the lab notify the ward or clinic immediately so that a new specimen may be collected? | If the Gram stain from the bottle is positive, does the lab call the result to the physician immediately? | When a positive blood culture broth is subcultured, is a chocolate plate included to ensure recovery of fastidious organisms? | Does the lab inoculate more than one patient sample on the same petri dish? | Does the lab perform AST on organisms that are possible contaminants? | Does the lab inoculate more than one patient sample on the same petri dish? | Does the lab inoculate more than one patient sample on the same petri dish? | When an ID result (bionumber) does not reach the threshold for an acceptable identification, is there evidence that appropriate action is taken, such as repeating the test by another method or performing additional biochemical tests? | Is catalase testing performed prior to coagulase testing on suspected Staph isolates? | Are negative slide coagulase results confirmed with a tube coagulase test before being reported? | If the Optochin result is equivocal (9-13mm), is bile solubility or other additional testing performed to confirm the ID? | When the instrument software flags an ID result as questionable, is there evidence that appropriate action is taken, such as repeating the test by another method or performing additional biochemical tests? | Does the lab ever intentionally mix two different organisms in the same inoculum for AST? | Is there evidence that appropriate actions are taken when the AST instrument software flags an AST result as questionable (such as checking for purity or repeating the test by another method)? | Is there evidence of such actions being taken? | Look at the cefotaxime disks currently in use. Does the drug concentration correspond correctly to the standard the lab uses? (CLSI breakpoints require 30ug disks, EUCAST breakpoints require 5ug disks). | Look at the ceftazidime disks currently in use. Does the drug concentration correspond correctly to the standard in use? (CLSI breakpoints require 30ug disks, EUCAST breakpoints require 10ug) | Look at the piperacillin-tazobactam disks currently in use. Does the drug concentration correspond correctly to the standard in use? (CLSI breakpoints require 100/10ug disks, EUCAST breakpoints require 30/6ug disks). | 1st generation cephalosporins (cefazolin, cephalothin, cephapririn, cephadrine) | 2nd generation cephalosporins (cefuroxime, cefonicid, cefamandole) | Cephamycins (cefoxitin, cefotetan) | Aminoglycosides (gentamicin, tobramycin, amikacin) | Does the lab use Nalidixic Acid to screen Salmonella isolates for ciprofloxacin resistance? | Does the lab use the correct fluoroquinolone (FQ) breakpoints for Salmonella spp? (Enterobacteriaceae FQ breakpoints should not be used for Salmonella spp). | Labs that do NOT use current cephalosporin and aztreonam breakpoints must perform routine ESBL phenotype testing. For ESBL-positive isolates, all penicillins, cephalosporins, and aztreonam that test susceptible must be reported as resistant. Is this practice (changing ESBL + interpretations from S to R) in place? | Labs that do NOT use current aztreonam and cephalosporin breakpoints should attach a warning comment to the report for ESBL positive organisms: “ESBL-producers should be considered clinically resistant to all penicillins, cephalosporins, and aztreonam.” Is this practice in place? | Labs that do NOT use current carbapenem breakpoints must perform routine testing for carbapenemase production (e.g. CarbaNP, mCIM, or a molecular assay). If a carbapenemase is detected, all carbapenems that test susceptible must be reported as resistant. Is this practice (changing results from S to R based on positive carbapenemase test result) in place? | Modified Hodge test | Disk diffusion | Gradient strip (e.g., Etest/Liofilchem) | Automated instrument (e.g., Vitek/Phoenix) | Broth microdilution (BMD) with Polysorbate 80 | Does the lab use Oxacillin disks to test for MRSA? | When oxacillin and cefoxitin results are discrepant for S. aureus (one is S and one is R), how does the lab report oxacillin? | Does the lab perform Staph aureus AST on any beta-lactam antibiotics other than penicillin, oxacillin, cefoxitin, or ceftaroline? | Does the lab use vancomycin disks to test for VISA/VRSA? | Are Staph aureus that are resistant to Erythromycin and susceptible or intermediate to Clindamycin tested for inducible clindamycin resistance? | Penicillin | Amoxicillin | Ampicillin | Cefotaxime | Ceftriaxone | Cefuroxime | Cefepime | Ertapenem | Meropenem | Imipenem | Are S.pneumoniae that are resistant to Erythromycin and susceptible or intermediate to Clindamycin tested for inducible clindamycin resistance? | When the ICR test is positive, is the clindamycin result changed to resistant? | 1st generation cephalosporins (cefazolin, cephalothin, cephapririn, cephadrine) | 2nd generation cephalosporins (cefuroxime, cefonicid, cefamandole) | Cephamycins (cefoxitin, cefotetan) | Clindamycin | Macrolides (Erythromycin, Azithromycin, Clarithromycin) | Tetracyclines (Tetracycline, Minocycline, Doxycycline) | Fluoroquinolones (Ciprofloxacin, Levofloxacin, Moxifloxacin) | Nitrofurantoin | If the lab practices cascade reporting, is it done in a way which ensures that the AST results excluded from the patient report are NOT excluded from the LIS or other main data repository? | Biosafety cabinets (Class IIA) | Covers on each centrifuge bucket | Cover over centrifuge rotor | Hand-washing station | Eyewash station/bottle | Sharps containers | Flame cabinet (for securely storing flammable liquids, e.g. ethanol) | Spill kit | First aid kit | Have all biosafety cabinets been recertified within a year of today’s date? | Gowns | Gloves | Goggles | Aerosol face protection (mask, face shield, or splatter guard) | Does lab policy require microbiology staff to wear close-toed shoes? | Is PPE utilized appropriately and consistently by laboratory staff? (Observe) | Does lab policy prohibit eating, drinking, and smoking in the laboratory? | Designated specifically for storage of media/reagents? | Free of staff food items? | Free of patient samples? | Well organized and free of clutter? | Are all hazardous chemicals stored appropriately (acids separate from alkaline; flammables in a flame cabinet)? | Is work area (bench and hood) disinfection documented daily? | Is there documentation demonstrating that an annual safety/biosafety refresher course is conducted for all staff handling specimens, isolates, or chemicals? | Is there documentation showing that the Quality Officers and Focal Points have received appropriate training in Quality Management Systems (QMS)? | Is there documentation showing that the the Supervisor/Quality Officer received training on how to effectively troubleshoot QC failures? | Is there evidence the supervisor or Quality Officer has received adequate training on how to perform root-cause analysis of QC failures? | Is there evidence the supervisor or Quality Officer has received adequate training on how to perform root-cause analysis for EQA failures? | Have technologists been adequately trained to recognize a poorly collected urine specimen (predominance of fecal or skin flora) based on the relative quantities, types, and mix of organisms present? | Is there evidence that microbiology staff have received adequate training to recognize intrinsic resistance patterns? (1: Yes - 2: Some, but would like additional training - 3: No) | Is there evidence that microbiology staff have received adequate training to recognize unusual or unexpected AST results that might require investigation? (e.g. Klebsiella spp. S to ampicillin; Staph spp. I/R to vancomycin) | Is there evidence that the supervisor recieved appropriate training on how to recognize unusual AST findings? | Is there evidence that microbiology staff have received adequate training on how to use the CLSI M100 or EUCAST documents effectively? (1: Yes - 2: Some, but would like additional training - 3: No) | Do laboratory staff understand the current limitations associated with colistin AST? (i.e., the risk of false susceptible results when using any method other than BMD without Polysorbate?) | In the last 6 months, has prolonged power failure disrupted the ability to provide routine bacteriology services? | Describe the internet service in the laboratory | In the last 6 months, has prolonged instrument failure disrupted the ability to provide routine bacteriology services? | In the last 6 months, have any stock outs disrupted the lab's ability to provide routine bacteriology services? | Does the laboratory use a computer-based Laboratory Information System (LIS)? | Does the hospital use a Hospital Information System (HIS) or Electronic Medical Record (EMR)? | Are inpatients assigned a unique patient ID number upon admission to the hospital? | Are outpatients assigned a unique patient ID number upon registration at the clinic? | Are patient ID numbers assigned in such a way that no two patients are given the same number in the course of one year? | Do patients retain the same patient ID number each time they are admitted to the hospital? | Do laboratory computers have antivirus software? | Do laboratory computers have genuine (not pirated) Operating Systems? | Does at least 50% of the technical staff possess formal education in microbiology or medical laboratory science? (Refer to the figure in column D) | Is the lab sufficiently staffed to provide high quality services? (Including support staff.) | How many times per year does the lab currently receive EQA/PT challenges that include both bacterial identification & AST? (Please do not include challenges designed to focus on a single organism, e.g., TB or N.gonorrhoeae) | On average, how long does the lab have to wait before receiving the results of their PT/EQA performance? | Is there internet in the lab to access free EUCAST PDFs or CLSI M100 online version? | Does lab policy primarily determine which isolates recieve AST, or is AST performed only when it is specifically requested by the doctor? | Does lab policy primarily determine which antibiotics to test and report, or does the lab only test and report the antibiotics specifically requested by the physician? | Does the hospital have an Antibiotic Stewardship Committee? | If the hospital has an Antibiotic Stewardship Committee, is a microbiolgist a member? | Does the hospital have a Pharmacy and Therapeutics Committee? | If the hospital has a Pharmacy and Therapeutics Committee, is a microbiolgist a member? | Does the hospital's Antibiotic Stewardship or Pharmacy and Therapeutic Committee meet at least annually to review national or international AST panel recommendations and modify them based on the hospital's formulary and cumulative antibiogram? | Does the lab prepare any media or distilled water? (e.g., blood agar, Mueller Hinton agar, blood culture bottles) | Does the lab have an automated blood culture instrument? (indicate manufacturer and model in comments) | Is the instrument functional today? | Is a service contract in place? | Is the software up to date? | Does the lab have an automated instrument for bacterial ID and AST? (e.g., Vitek, Microscan, Phoenix) | Is the instrument functional today? | Is a service contract in place? | Is the software up to date? | Does the lab have an automated instrument for reading disk diffusion? (e.g. SIRSCAN, BIOMIC V3, ADAGIO, etc.) | Is the instrument functional today? | Is a service contract in place? | Is the software up to date? | Does the lab have a MALDI instrument for organism ID? (e.g., Bruker, Biomerieux) | Is the instrument functional today? | Is a service contract in place? | Is the software up to date? | Does the lab have a PCR instrument used for detecting antibiotic resistance genes? (e.g., GeneXpert) | Is the instrument functional today? | Is a service contract in place? | Is the software up to date? | Patient Last Name/Surname | Patient First Name | Patient Identification Number | Patient Date of Birth | Patient Age | Patient Gender | Patient Location (Ward or Unit at the time of specimen collection, e.g., "ICU") | Patient Date of Admission | Specimen identification number | Specimen Type (e.g. Wound) | Specimen Source/Body Site (e.g. Arm) | Additional descriptors (e.g., Left, Right) | Date of specimen collection | Time of specimen collection | Date of specimen receipt | Time of specimen receipt | Quantity of Epithelial Cells per low power field | Quantity of PMNs (WBCs) per low power field | Quantity of bacterial cells per high power field | Type of bacterial cells (gram positive cocci, gram negative bacilli, etc.) | Description of colony morphologies (e.g. "mucoid lactose fermenter" or "beta-hemolytic") | Description of colony quantities (e.g. "1+, 2+, 3+, 4+" or "few, moderate, many") | Gram stain of bacterial colony | Biochemical test results (e.g., "catalase positive") for conventional test methods | Organism name | Isolate number (e.g., when more than one pathogen is encountered in a culture: isolate #1, isolate #2) | Can the LIS record the AST method used to obtain each individual antibiotic result (e.g., Etest vs. Vitek vs. disk)? | Disk diffusion zone sizes | Disk diffusion interpretation (S/I/R) | MIC values | MIC interpretation (S/I/R) | Can the LIS record MIC values to three decimal places (e.g. 0.016)? | Can the LIS suppress (hide) an individual antibiotic result from the patient report without deleting it from the database (for cascade reporting)? | Does the LIS software automatically interpret zone sizes into Susceptible, Intermediate, Resistant? | Does the LIS software automatically interpret MICs into Susceptible, Intermediate, Resistant? | If the LIS software automatically interprets zone sizes or MICs, are the breakpoints updated annually? | If the LIS software automatically interprets zone sizes or MICs, are the breakpoints up to date today? | Can the LIS de-duplicate data based on select criteria (e.g., patient ID, organism, specimen date)? | Can the LIS produce a cumulative antibiogram report? | Can the LIS interface with automated AST instruments (e.g., Vitek, Phoenix, SIRScan, BIOMIC)? | Can the LIS interface with the Hospital Information System (HIS)? | Can the LIS export line lists of data to .txt or .csv files? | If the lab uses an automated AST instrument, describe the data flow between the LIS and the instrument software. | Does the hospital use a Hospital Information System (HIS) or Electronic Medical Record (EMR)? | If the LIS and HIS/EMR are interfaced, describe the data flow between the LIS and the HIS/EMR | What is the source of the blood used to make the blood agar, chocolate, and/or MHB plates? | Do records demonstrate that QC is performed on each newly reconstituted batch or newly received lot number/shipment of media? | Do QC records for blood agar plates demonstrate that they are checked for their ability to support growth of fastidious organisms such as Streptococcus pneumoniae? | Do QC records for blood agar plates demonstrate that they are checked for their ability to show beta, alpha, and gamma hemolysis? | Do QC records for chocolate agar plates demonstrate that they are checked for their ability to support the growth of fastidious organisms, such as Neisseria gonorrhoeae or H.influenzae? | MacConkey (MAC) and Eosin methylene blue (EMB) agars contain bile salts and/or dyes that are toxic for Gram positive bacteria when made properly. Do QC records for MAC and/or EMB plates demonstrate that each batch/lot is challenged using a Gram positive organism? | Dyes and pH indicators in MAC and EMB plates provide a color indicator to distinguish between lactose fermenting (LF) and non-lactose fermenting (NLF) Gram negative organisms. Do QC records for MAC and/or EMB plates demonstrate that each batch/lot is challenged using both LF and NLF organisms? | Do QC records for selective stool agar plates (e.g. XLD, SS, HE) demonstrate that they are checked for their ability to suppress the growth of Gram positive organisms? | Do QC records for selective stool agar plates demonstrate that they are checked for their ability to make H2S (hydrogen sulfide) production visible using an H2S producing organism, such as Salmonella spp or Proteus vulgaris? | Do QC records for selective stool agar plates demonstrate that they are checked for their ability to make the acid byproducts of carbohydrate fermentation visible using both fermenters and nonfermenters? | Pseudomonas aeruginosa 27853 and gentamicin disk | Enterococcus faecalis 29212 or 33186 and trimethoprim-sulfamethoxazole disk | Does the lab prepare blood culture bottles in-house? | Following manufacturer instructions, are all of the recommended ATCC strains in use for the identification kits? | Following manufacturer instructions, are all of the recommended ATCC strains in use for the automated instrument ID cards/trays? | Which blood culture incubation systems does the lab use? | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | What is the source of the plasma used for coagulase testing? | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this reagent in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this test in use in your lab? (If No, select N/A for the remaining questions about this reagent) | Is this test in use in your lab? (If No, select N/A for the remaining questions about this reagent) | When the primary plate has mixed colony types, is it standard practice to subculture each colony of interest to a fresh plate to ensure purity prior to pursuing identification? | Is it standard practice to perform a gram stain on each isolate of interest prior to performing any other testing? | For gram negative bacilli, is it standard practice to perform an oxidase test first, before proceeding with any other identification tests (including automated ID)? | For gram negative bacilli, is it standard practice to perform an indole test second, before proceeding with other identification tests (including automated ID)? | For oxidase-negative gram negative bacilli that do not ferment lactose (clear on MacConkey), are sufficient tests available to achieve a definitive identification? | For oxidase-positive gram negative bacilli that are not Pseudomonas aeruginosa (lack the characteristic odor), are sufficient tests available to achieve a definitive identification? | For gram positive cocci, is it standard practice to perform a catalase test first, before proceeding with any other identification tests (including automated ID)? | For catalase positive gram positive cocci, is it standard practice to perform a coagulase test next, before proceeding with other identification tests (including automated ID)? | For catalase negative gram positive cocci, is it standard practice to evaluate the type of hemolysis (alpha, beta, gamma), before proceeding with other identification tests (including automated ID)? | Are the in-use antibiotic disks and strips stored in a tightly sealed container with active desiccants? | Do the desiccants change color as moisture levels increase (indicating the need to replace or recharge)? | If desiccants do not have a color indicator, are colorless desiccants replaced least monthly? | After AST inoculation, are “purity plates” made from the remaining suspension? | Is the automated AST instrument software up to date? | Do the AST SOPs or bench aides provide examples of intrinsic resistance patterns? (Such as those found in CLSI M100 Appendix B or EUCAST Expert Rules V3.1) | Do the AST SOPs or bench aides define examples of unusual or unexpected AST results? (Such as those found in CLSI M100 Appendix A or EUCAST Expert Rules V3.1) | Do the AST SOPs or bench aides describe what actions to take when unusual or unexpected AST results are encountered (e.g., check purity, reconfirm organism ID, check relevant QC, repeat testing, notify supervisor)? | Does a supervisor review all AST results for unusual findings before results are given to physicians ? | 1: CLSI - 2: EUCAST - 3: Other (please list in comments) - 4: None/mixed | Ask to see the lab’s most current hard copy of the standard. Is it less than 3 years old? | Does the lab obtain updates of the standard in use at least every 3 years? | Does the lab review important standards changes, eg. breakpoint changes, with the relevant hospital committees (e.g. pharmacy and therapeutics, stewardship)? | For labs that DO use current cephalosporin and aztreonam breakpoints, CLSI and EUCAST no longer recommend routine testing for ESBL phenotype. Furthermore, if ESBL testing is performed and the test is positive, interpretations for beta-lactam agents do NOT need to be changed from susceptible to resistant. Has the lab discontinued editing AST results based on the ESBL result? | Does the lab perform any genotypic tests for ESBL production? (E.g. PCR) | Do records indicate that quality control for ESBL testing is done either on a weekly basis or each time the test is performed? | Do records indicate that lab uses both positive and negative control organisms to QC the ESBL test in use? (A commonly used ESBL positive strain is Klebsiella pneumoniae ATCC 700603) | For labs that DO use current carbapenem breakpoints, CLSI and EUCAST no longer recommend routine testing for carbapenemase production. Furthermore, if such testing is performed and the test is positive, interpretations for carbapenems do NOT need to be changed from susceptible to resistant. Has the lab discontinued editing AST results based on the carbapenemase result? | Other disk method, e.g., combination disk test or double disk synergy | MIC Strip test, e.g., Etest KPC, MBL or Liofilchem MRP/MBO, ETP/EBO | Biochemical (colorimetric) test, e.g. CarbaNP, BCT, or β CARBA | Chromogenic agar specific for carbapenemase producers | Modified carbapenem inactivation method (MCIM) | Does the lab perform any genotypic tests for carbapenemase production? (E.g. PCR, GeneXpert, etc.) | Do records indicate that quality control is done each time carbapenemase testing is performed? | Do records indicate that lab uses both positive and negative control organisms to QC the carbapenemase test in use? | Broth microdilution (BMD) without Polysorbate 80 | Does the lab perform colistin AST? | Colistin Broth Disk Elution (CBDE) Test | Do records indicate that quality control for colistin AST is performed on either a weekly basis or each time the test is performed? | Do records indicate that lab uses both susceptible and resistant organisms to QC the colistin test in use? (E.coli NCTC 13846 is recommended to detect low level resistance) | When colistin resistance is detected, is the isolate sent to a reference lab for molecular characterization (e.g., testing for mcr genes)? | If the lab uses broth microdilution for colistin AST, is colistin sulfate used, not colistin methane sulfonate (sulfomethate)? | Does the lab test Staph aureusisolates against penicillin? | Are Staph aureus isolates with penicillin zones sizes or MICs in the susceptible range tested for β-lactamase production using the zone-edge test before being reported as penicillin susceptible? | When a vancomycin MIC >8 is detected for Staph aureus, is the isolate sent to a referral lab for confirmation testing and further characterization? | Does the lab perform AST for S.pneumoniae? (Not scored. If No, skip to next section.) | Penicillin | Ceftriaxone and/or Cefotaxime | When S.pneumoniae is isolated from CSF, are penicillin, ceftriaxone, and/or cefotaxime reported using the meningitis breakpoints only? | When S.pneumoniae is isolated from specimens other than CSF, are penicillin, ceftriaxone, and/or cefotaxime reported using both meningitis and non-meningitis breakpoints? | Does the lab perform the test for Inducible Clindamycin Resistance (ICR), also known as the “D-test” on Staph aureusand/or Strep pneumoniae? | Does the SOP for the ICR test specify that the erythromycin and clindamycin disks must be placed 15-26 mm apart for Staph species? | Does the SOP for the ICR test specify that the erythromycin and clindamycin disks must be placed 12 mm apart for Strep species? | Do records indicate that quality control for ICR testing is done either on a weekly basis or each time the test is performed? | Do records indicate that lab uses both positive and negative control organisms to QC the ICR test in use? (Commonly used ICR positive strain is S.aureus ATCC BAA-977) | Does the lab produce a cumulative antibiogram at least annually? | Does the lab practice “cascade reporting”?
# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184 | 185 | 186 | 187 | 188 | 189 | 190 | 191 | 192 | 193 | 194 | 195 | 196 | 197 | 198 | 199 | 200 | 201 | 202 | 203 | 204 | 205 | 206 | 207 | 208 | 209 | 210 | 211 | 212 | 213 | 214 | 215 | 216 | 217 | 218 | 219 | 220 | 221 | 222 | 223 | 224 | 225 | 226 | 227 | 228 | 229 | 230 | 231 | 232 | 233 | 234 | 235 | 236 | 237 | 238 | 239 | 240 | 241 | 242 | 243 | 244 | 245 | 246 | 247 | 248 | 249 | 250 | 251 | 252 | 253 | 254 | 255 | 256 | 257 | 258 | 259 | 260 | 261 | 262 | 263 | 264 | 265 | 266 | 267 | 268 | 269 | 270 | 271 | 272 | 273 | 274 | 275 | 276 | 277 | 278 | 279 | 280 | 281 | 282 | 283 | 284 | 285 | 286 | 287 | 288 | 289 | 290 | 291 | 292 | 293 | 294 | 295 | 296 | 297 | 298 | 299 | 300 | 301 | 302 | 303 | 304 | 305 | 306 | 307 | 308 | 309 | 310 | 311 | 312 | 313 | 314 | 315 | 316 | 317 | 318 | 319 | 320 | 321 | 322 | 323 | 324 | 325 | 326 | 327 | 328 | 329 | 330 | 331 | 332 | 333 | 334 | 335 | 336 | 337 | 338 | 339 | 340 | 341 | 342 | 343 | 344 | 345 | 346 | 347 | 348 | 349 | 350 | 351 | 352 | 353 | 354 | 355 | 356 | 357 | 358 | 359 | 360 | 361 | 362 | 363 | 364 | 365 | 366 | 367 | 368 | 369 | 370 | 371 | 372 | 373 | 374 | 375 | 376 | 377 | 378 | 379 | 380 | 381 | 382 | 383 | 384 | 385 | 386 | 387 | 388 | 389 | 390 | 391 | 392 | 393 | 394 | 395 | 396 | 397 | 398 | 399 | 400 | 401 | 402 | 403 | 404 | 405 | 406 | 407 | 408 | 409 | 410 | 411 | 412 | 413 | 414 | 415 | 416 | 417 | 418 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 426 | 427 | 428 | 429 | 430 | 431 | 432 | 433 | 434 | 435 | 436 | 437 | 438 | 439 | 440 | 441 | 442 | 443 | 444 | 445 | 446 | 447 | 448 | 449 | 450 | 451 | 452 | 453 | 454 | 455 | 456 | 457 | 458 | 459 | 460 | 461 | 462 | 463 | 464 | 465 | 466 | 467 | 468 | 469 | 470 | 471 | 472 | 473 | 474 | 475 | 476 | 477 | 478 | 479 | 480 | 481 | 482 | 483 | 484 | 485 | 486 | 487 | 488 | 489 | 490 | 491 | 492 | 493 | 494 | 495 | 496 | 497 | 498 | 499 | 500 | 501 | 502 | 503 | 504 | 505 | 506 | 507 | 508 | 509 | 510 | 511 | 512 | 513 | 514 | 515 | 516 | 517 | 518 | 519 | 520 | 521 | 522 | 523 | 524 | 525 | 526 | 527 | 528 | 529 | 530 | 531 | 532 | 533 | 534 | 535 | 536 | 537 | 538 | 539 | 540 | 541 | 542 | 543 | 544 | 545 | 546 | 547 | 548 | 549 | 550 | 551 | 552 | 553 | 554 | 555 | 556 | 557 | 558 | 559 | 560 | 561 | 562 | 563 | 564 | 565 | 566 | 567 | 568 | 569 | 570 | 571 | 572 | 573 | 574 | 575 | 576 | 577 | 578 | 579 | 580 | 581 | 582 | 583 | 584 | 585 | 586
Code | Name | ADD | City | PROV | DIS | Country | ASS1 | ASS2 | ASS3 | Date | Z | Y | X | Fund | type | level | Slev | Beds | DD | Grad | bmicro | bmacro | AgarD | Vit | Pho | Mic | Other | OESBL | OCRE | OMRSA | OVRE | OCOLI | OOTH | GESBL | GCARBA | GMECA | GVAN | GMRC | GOTH | SLIP | DSLIP | SSLIP | LQSI | LQSIP | OMENT | ABLOOD | ASTOL | AURINE | AORGA | AAST | AOTH | BODY | WLBLOOD | WLURIN | WLSTOOL | WLSS | WLVC | WLYE | WLCJ | WLO157 | WLRESPI | WLWOUND | WLCSF | WLBF | WFGENIT | WLANAERO | WLYEAST | WLMOLD | WLMRSA | WLVRE | WLCRE | WLREFER | WLOTH | AAST | DD | GS | BMIC | BMAC | AGAD | BIQC1 | BIQC2 | BIQC3 | BIQC4 | BIQC5 | BIQC6 | BIQC7 | BIQC8 | BIQC9 | BIQC10 | BIQC11 | BIQC12 | BIQC13 | BIQC14 | BIQC15 | BIQC16 | BIQC17 | BIQC18 | BIQC19 | BIQC20 | BIQC21 | BIQC22 | BIQC23 | BIQC24 | BIQC25 | BISOP1 | BISOP2 | BISOP3 | BISOP4 | BISOP5 | BISOP6 | BISOP7 | BISOP8 | BISOP9 | BISOP10 | BISOP11 | BISOP12 | BISOP13 | BISOP14 | BISOP15 | BISOP16 | BISOP17 | BISOP18 | BISOP19 | BISOP20 | BISOP21 | BISOP22 | BISOP23 | BISOP24 | BISOP25 | ADVMIC | ADVOTH | PGMIC | PGOTH | BMIC | BOTH | UGMIC | UGOTH | HS | OTJ | SOTH | %MED | %GRA | MCF | RULCAL | BB | CL | DENS | PIP | CENT | MICRO | THERM | CO2 | CJ | INC | REF | F-20 | F-60 | F-80 | DESS | HAO | BSCIIA | CAUTO | DAUTO | PH | BAL | CONDUC | WD | HPMS | WB | BCI | ISAST | DDR | MT | PCRAST | RF | TO | SF | GI | FAC | LIS | DATA | QA | MPREP | IDQC | ASTQC | SPEC | PROC | ID | BAST | EAST | PANE | SAFE | FAC1 | FAC2 | FAC3 | FAC4 | FAC5 | FAC6 | FAC7 | FAC8 | FAC9 | FAC10 | FAC11 | LIS1 | LIS2 | LIS3 | LIS4 | LIS5 | LIS6 | DATA1 | DATA2 | DATA3 | DATA4 | DATA5 | DATA6 | DATA7 | QA1 | QA2 | QA3 | QA4 | MPREP1 | MPREP2 | MPREP3 | MPREP4 | MPREP5 | MPREP6 | IDQC1 | IDQC2 | IDQC3 | IDQC4 | IDQC5 | IDQC6 | IDQC7 | IDQC8 | ASTQC1 | ASTQC2 | ASTQC3 | ASTQC4 | ASTQC5 | SPEC1 | SPEC2 | SPEC3 | SPEC4 | SPEC5 | PROC1 | PROC2 | PROC3 | PROC4 | ID1 | ID2 | ID3 | ID4 | ID5 | ID6 | ID7 | ID8 | ID9 | ID10 | ID11 | ID12 | ID13 | ID14 | ID15 | BAST1 | BAST2 | BAST3 | BAST4 | BAST5 | BAST6 | EAST1 | EAST2 | EAST3 | EAST4 | EAST5 | EAST6 | EAST7 | EAST8 | EAST9 | EAST10 | PAN1 | PAN2 | PAN3 | SAFE1 | SAFE2 | SAFE3 | SAFE4 | RF1 | RF2 | RF3 | RF4 | RF5 | RF6 | RF7 | RF8 | RF9 | RF10 | RF11 | RF12 | RF13 | RF14 | RF15 | RF16 | RF17 | RF18 | RF19 | RF20 | RF21 | RF22 | RF23 | RF24 | RF25 | RF26 | RF27 | RF28 | RF29 | RF30 | RF31 | RF32 | RF33 | RF34 | RF35 | RF36 | RF37 | RF38 | RF39 | RF40 | RF41 | RF42 | RF43 | RF44 | RF45 | RF46 | RF47 | RF48 | RF49 | RF50 | RF51 | RF52 | RF53 | RF54 | RF55 | RF56 | RF57 | RF58 | RF59 | RF60 | RF61 | RF62 | RF63 | RF64 | RF65 | RF66 | RF67 | RF68 | RF69 | RF70 | RF71 | RF72 | RF73 | RF74 | RF75 | RF76 | RF77 | RF78 | RF79 | RF80 | RF81 | RF82 | RF83 | RF84 | RF85 | RF86 | RF87 | RF88 | RF89 | RF90 | RF91 | RF92 | RF93 | RF94 | RF95 | RF96 | RF97 | RF98 | RF99 | RF100 | RF101 | TO1 | TO2 | TO3 | TO4 | TO5 | TO6 | TO7 | TO8 | TO9 | TO10 | SF1 | SF2 | SF3 | SF4 | SF5 | SF6 | SF7 | SF8 | SF9 | SF10 | SF11 | SF12 | SF13 | SF14 | SF15 | SF16 | SF17 | SF18 | SF19 | SF20 | SF21 | SF22 | SF23 | SF24 | Q1.1 | Q1.2 | Q1.3 | Q1.4 | Q1.5 | Q1.6 | Q1.7 | Q1.8 | Q1.9 | Q1.10 | Q1.11 | Q1.12 | Q1.13 | Q1.14 | Q1.15 | Q1.16 | Q1.17 | Q1.18 | Q1.19 | Q1.20 | Q1.21 | Q2.1 | Q2.2 | Q2.3 | Q2.4 | Q2.5 | Q2.6 | Q2.7 | Q2.8 | Q2.9 | Q2.10 | Q2.11 | Q2.12 | Q2.13 | Q2.14 | Q2.15 | Q2.16 | Q2.17 | Q2.18 | Q2.19 | Q2.20 | Q2.21 | Q2.22 | Q2.23 | Q2.24 | Q2.25 | Q2.26 | Q2.27 | Q2.28 | Q2.29 | Q2.30 | Q2.31 | Q2.32 | Q2.33 | Q2.34 | Q2.35 | Q2.36 | Q2.37 | Q2.38 | Q2.39 | Q2.40 | Q2.41 | Q2.42 | Q2.43 | Q2.44 | Q2.45 | Q5.1 | Q5.2 | Q5.3 | Q5.4 | Q5.5 | Q5.6 | Q5.7 | Q5.8 | Q5.9 | Q5.10 | Q5.11 | Q5.12 | Q5.13 | Q6.1 | Q6.2 | Q9.1 | Q10.1 | Q10.2 | Q10.3 | Q10.4 | Q10.5 | Q10.6 | Q10.7 | Q10.8 | Q10.9 | Q10.10 | Q10.11 | Q10.12 | Q10.13 | Q10.14 | Q10.15 | Q10.16 | Q10.17 | Q10.18 | Q10.19 | Q10.20 | Q10.21 | Q10.22 | Q10.23 | Q10.24 | Q10.25 | Q10.26 | Q10.27 | Q10.28 | Q10.29 | Q10.30 | Q10.31 | Q10.32 | Q10.33 | Q10.34 | Q10.35 | Q11.1 | Q11.2 | Q11.3 | Q11.4 | Q11.5 | Q11.6 | Q11.7 | Q11.8 | Q11.9 | Q11.10 | Q11.11 | Q11.12 | Q11.13 | Q12.1 | Q12.2 | Q12.3 | Q12.4 | Q12.5 | Q12.6 | Q12.7 | Q12.8 | Q12.9 | Q12.10 | Q12.11 | Q12.12 | Q12.13 | Q12.14 | Q12.15 | Q12.16 | Q12.17 | Q12.18 | Q12.19 | Q12.20 | Q12.21 | Q12.22 | Q12.23 | Q12.24 | Q12.25 | Q12.26 | Q12.27 | Q12.28 | Q12.29 | Q12.30 | Q12.31 | Q12.32 | Q12.33 | Q13.1 | Q13.2