## LIVD Publication
Publisher | APHL
LOINC Copyright | This material contains content from LOINC® (http://loinc.org). The LOINC table, LOINC codes, LOINC panels and form file, LOINC linguistic variants file, LOINC/RSNA Radiology Playbook, and LOINC/IEEE Medical Device Code Mapping Table are copyright © 1995-2022, Regenstrief Institute, Inc. and the Logical Observation Identifiers Names and Codes (LOINC) Committee and is available at no cost under the license at http://loinc.org/terms-of-use.
LOCALIZATION | en-US
Document Identifier | Lyme_LIVD
Publication Date | 2024-07-03
Publisher URL | https://www.cdc.gov/csels/dls/livd-codes.html
Publisher Statement | This is the offically published LIVD file for Lyme disease diagnostic tests that are FDA approved (EUAs or 510k approval). If you have questions about this mapping, please send email to informatics.support@aphl.org
SNOMED CT Copyright | This material includes SNOMED Clinical Terms® (SNOMED CT®) which is used by permission of the International Health Terminology Standards Development Organisation (IHTSDO). All rights reserved. SNOMED CT®, was originally created by The College of American Pathologists. “SNOMED” and “SNOMED CT” are registered trademarks of the IHTSDO.
Accessibility | If you need assistance to read the Excel spreadsheet presented here, please email informatics.support@aphl.org and use “Lyme LIVD Test Code Mapping” in the subject line.

## LOINC Mapping
Master Row ID | Lab Test ID | Manufacturer | Model | Vendor Analyte Name | Vendor Specimen Description | Vendor Result Description | Test Performed LOINC Code | Test Performed LOINC Long Name | Test Ordered LOINC Code | Test Ordered LOINC Long Name | Vendor Comment | Vendor Analyte Code | Vendor Reference ID | Testkit Name ID | Testkit Name ID Type | Equipment UID | Equipment UID Type | Component | Property | Time | System | Scale | Method | Publication Version ID | LOINC Version ID | Authorized Settings | Test Type | Test Method | Notes
|  | ViraMed Biotech AG | Borrelia All-In-One ViraChipTest Kit | Borrelia All-In-One ViraChip® Result | Serum (119364003^Serum specimen^SCT) | Positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) Negative (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT)) invalid (455371000124106^Invalid result^SCT or 125154007^Specimen unsatisfactory for evaluation^SCT) | 101353-1 | Borrelia burgdorferi Ab [Interpretation] in Serum Qualitative by IA.MTTT | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel |  |  | V-BTCFOK / V-BTCFDK | ViraMed Biotech AG_Borrelia All-In-One ViraChip Test Kit | MNT | ViraMed Biotech AG_ViraChip Reader | MNI | Borrelia burgdorferi Ab.IgG & IgM MTT Algorithm | Imp | Pt | Ser | Ord | IA |  | 2.74 | High Complexity |  |  | this chip tests for B burgdor VlsE1 IgG + IgM AB and if any are detected then reflexes automatically to B burgdor IgG+IgM and provides the interpretation of this algorithm
|  | Zeus Scientific, Inc | Borrelia VlsE1/pepC10 IgG/IgM Test System | Zeus Lyme VlsE1 and pepC10 IgG/IgM Ab EIA | Serum (119364003^Serum specimen^SCT) | Index Value/OD Ratio Negative Specimens ≤0.90 (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Equivocal Specimens 0.91 to 1.09 (42425007^Equivocal^SCT) Positive Specimens ≥1.10 (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 101403-4 | Borrelia burgdorferi.VlsE1+pepC10 IgG+IgM Ab [Presence] in Serum by Immunoassay | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel |  |  |  | B. burgdorferi IgG Test System_Zeus Scientific, Inc | MNT | <no specific equipment - Use of a single (450nm), or dual (450/620 - 650nm), wavelength reader is acceptable.> | N/A | Borrelia burgdorferi.VlsE1+pepC10 IgG+IgM Ab | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | Zeus Scientific, Inc | Borrelia VlsE1/pepC10 IgG/IgM Test System | Lyme IgG/IgM Ab EIA | Serum (119364003^Serum specimen^SCT) | Index Value/OD Ratio Negative Specimens ≤0.90 (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Equivocal Specimens 0.91 to 1.09 (42425007^Equivocal^SCT) Positive Specimens ≥1.10 (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 83081-0 | Borrelia burgdorferi IgG+IgM Ab [Presence] in Serum by Immunoassay | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel |  |  |  | B. burgdorferi IgG Test System_Zeus Scientific, Inc | MNT | <no specific equipment - Use of a single (450nm), or dual (450/620 - 650nm), wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgG+IgM | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | Zeus Scientific, Inc | Borrelia VlsE1/pepC10 IgG/IgM Test System | Lyme IgM Ab EIA | Serum (119364003^Serum specimen^SCT) | Index Value/OD Ratio Negative Specimens ≤0.90 (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Equivocal Specimens 0.91 to 1.09 (42425007^Equivocal^SCT) Positive Specimens ≥1.10 (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel |  |  |  | B. burgdorferi IgG Test System_Zeus Scientific, Inc | MNT | <no specific equipment - Use of a single (450nm), or dual (450/620 - 650nm), wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgM | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.74 | High Complexity
|  | Zeus Scientific, Inc | Borrelia VlsE1/pepC10 IgG/IgM Test System | Lyme IgG Ab EIA | Serum (119364003^Serum specimen^SCT) | Index Value/OD Ratio Negative Specimens ≤0.90 (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Equivocal Specimens 0.91 to 1.09 (42425007^Equivocal^SCT) Positive Specimens ≥1.10 (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel |  |  |  | B. burgdorferi IgG Test System_Zeus Scientific, Inc | MNT | <no specific equipment - Use of a single (450nm), or dual (450/620 - 650nm), wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgG | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | Zeus Scientific, Inc | Borrelia VlsE1/pepC10 IgG/IgM Test System | Overall Lyme Exposure | Serum (119364003^Serum specimen^SCT) | supportive evidence for the presence of IgG/IgM antibodies and exposure to B. burgdorferi (260373001^Detected^SCT and send "supportive evidence for the presence of IgG/IgM antibodies and exposure to B. burgdorferi" as NTE after this OBX, unless this is already sent in OBX-5.9) no supportive evidence for the presence of IgG/IgM antibodies and exposure to B. burgdorferi (260415000^Not detected^SCT and send "no supportive evidence for the presence of IgG/IgM antibodies and exposure to B. burgdorferi" as NTE after this OBX, unless this is already sent in OBX-5.9) | 101353-1 | Borrelia burgdorferi Ab [Interpretation] in Serum Qualitative by IA.MTTT | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel | [7 empty] | Borrelia burgdorferi Ab.IgG & IgM MTT Algorithm | Imp | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | DiaSorin | LIAISON® Lyme Total Antibody Plus | Lyme Total Ab CLIA | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Index = < 0.90 = Negative = Absence of detectable Borrelia burgdorferi antibodies. A negative result does not exclude the possibility of Borrelia burgdorferi infection. If early Lyme disease is suspected, a second sample should be collected and tested two to four weeks later. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Index = ≥ 0.90 and < 1.10 = Equivocal = The imprecision inherent in this method does not allow definitive categorization of samples that read close to the cutoff. All equivocal samples should be tested further following the Standard Two-Tier Testing (STTT) or Modified Two-Tier Testing (MTTT) Methodologies. (42425007^Equivocal^SCT) Index = ≥ 1.10 = Positive = Presence of detectable Borrelia burgdorferi antibodies. All positive results should be tested further following the the Standard Two-Tier Testing (STTT) or Modified Two-Tier Testing (MTTT) Methodologies. (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 101354-9 | Borrelia afzelii+burgdorferi+garinii IgG+IgM Ab [Presence] in Serum by Immunoassay | 101404-2 | Borrelia afzelii+burgdorferi+garinii Ab.IgG and IgM panel - Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme Total Antibody Plus_DiaSorin Inc | MNT | LIAISON XL Analyzer_DiaSorin | MNI | Borrelia afzelii+burgdorferi+garinii Ab.IgG+IgM | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | DiaSorin | LIAISON® Lyme Total Antibody Plus | Lyme Total Ab CLIA | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Index = < 0.90 = Negative = Absence of detectable Borrelia burgdorferi antibodies. A negative result does not exclude the possibility of Borrelia burgdorferi infection. If early Lyme disease is suspected, a second sample should be collected and tested two to four weeks later. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Index = ≥ 0.90 and < 1.10 = Equivocal = The imprecision inherent in this method does not allow definitive categorization of samples that read close to the cutoff. All equivocal samples should be tested further following the Standard Two-Tier Testing (STTT) or Modified Two-Tier Testing (MTTT) Methodologies. (42425007^Equivocal^SCT) Index = ≥ 1.10 = Positive = Presence of detectable Borrelia burgdorferi antibodies. All positive results should be tested further following the the Standard Two-Tier Testing (STTT) or Modified Two-Tier Testing (MTTT) Methodologies. (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 101354-9 | Borrelia afzelii+burgdorferi+garinii IgG+IgM Ab [Presence] in Serum by Immunoassay | 101404-2 | Borrelia afzelii+burgdorferi+garinii Ab.IgG and IgM panel - Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme Total Antibody Plus_DiaSorin Inc | MNT | LIAISON XS Analyzer_DiaSorin | MNI | Borrelia afzelii+burgdorferi+garinii Ab.IgG+IgM | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | DiaSorin | LIAISON® Borrelia IgG | Lyme IgG Ab CLIA [AU] | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | <numeric value > in AU/mL The analyzer automatically calculates Borrelia burgdorferi IgG antibody concentrations expressed as arbitrary units (AU/mL) and grades the results - Assay range. 5 to 240 AU/mL Borrelia burgdorferi IgG. | 92810-1 | Borrelia afzelii+burgdorferi+garinii IgG Ab [Units/volume] in Serum by Immunoassay | 101404-2 | Borrelia afzelii+burgdorferi+garinii Ab.IgG and IgM panel - Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme Total Antibody Plus_DiaSorin Inc | MNT | LIAISON XL Analyzer_DiaSorin | MNI | Borrelia afzelii+burgdorferi+garinii Ab.IgG | Acnc | Pt | Ser | Qn | IA |  | 2.74 | High Complexity
|  | DiaSorin | LIAISON® Borrelia IgG | Lyme IgG Ab CLIA | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Samples with Borrelia burgdorferi IgG concentrations below 10 AU/mL should be graded negative. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Samples with Borrelia burgdorferi IgG concentrations ranging between 10 and 15 AU/mL should be graded equivocal. Equivocal samples must be retested in order to confirm the initial result. Samples which are positive at the second test should be considered positive. (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) Samples which are negative at the second test should be considered negative. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) A second sample should be collected and tested no less than one week later when the result is repeatedly equivocal. (42425007^Equivocal^SCT) Samples with Borrelia burgdorferi IgG concentrations equal to or above 15 AU/mL should be graded positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 92814-3 | Borrelia afzelii+burgdorferi+garinii IgG Ab [Presence] in Serum by Immunoassay | 101404-2 | Borrelia afzelii+burgdorferi+garinii Ab.IgG and IgM panel - Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme Total Antibody Plus_DiaSorin Inc | MNT | LIAISON XL Analyzer_DiaSorin | MNI | Borrelia afzelii+burgdorferi+garinii Ab.IgG | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | DiaSorin | LIAISON® Borrelia IgM II | Lyme IgM Ab CLIA | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Samples with Borrelia burgdorferi IgM levels below an index value of 0.9 should be graded negative. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Samples with Borrelia burgdorferi IgM levels ranging between an index value of 0.9 and 1.1 should be graded equivocal. Equivocal samples must be retested in order to confirm the initial result. Samples which are positive at the second test should be considered positive.(10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) Samples which are negative at the second test should be considered negative. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) A second sample should be collected and tested no less than one week later when the result is repeatedly equivocal. (42425007^Equivocal^SCT) Samples with Borrelia burgdorferi IgM levels equal to or above an index value of 1.1 should be graded positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 101355-6 | Borrelia afzelii+burgdorferi+garinii IgM Ab [Presence] in Serum by Immunoassay | 101404-2 | Borrelia afzelii+burgdorferi+garinii Ab.IgG and IgM panel - Serum or Plasma by Immunoassay |  |  |  | LIAISON Borrelia IgM II_DiaSorin Inc | MNT | LIAISON XL Analyzer_DiaSorin | MNI | Borrelia afzelii+burgdorferi+garinii Ab.IgM | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | DiaSorin | MTTT Interpretation | Overall Lyme Exposure | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | IgG- and IgM - (260415000^Not detected^SCT) - the clinical interpretation is: "No evidence of infection. In case of clinical uncertainty (presence of tick bite or neurological symptoms), the patients should be followed up during time." - this should be sent as NTE after this OBX IgG – and IgM + (711621000124102^Borrelia burgdorferi immunoglobulin M antibody detected^SCT) - the clinical interpretation is: "Probable infection at an early stage." - this should be sent as NTE after this OBX IgG+ and IgM- (711631000124104^Borrelia burgdorferi immunoglobulin G antibody detected^SCT) - the clinical interpretation is: "Probable infection at any stage." - this should be sent as NTE after this OBX IgG+ and IgM+ - (711611000124105^Borrelia burgdorferi immunoglobulin G and M antibodies detected^SCT ) - the clinical interpretation is: "Probable acute infection." - this should be sent as NTE after this OBX | 101356-4 | Borrelia afzelii+burgdorferi+garinii Ab [Interpretation] in Serum by IA.MTTT | 101404-2 | Borrelia afzelii+burgdorferi+garinii Ab.IgG and IgM panel - Serum or Plasma by Immunoassay | [7 empty] | Borrelia afzelii+burgdorferi+garinii Ab.IgG & IgM | Imp | Pt | Ser | Nom | IA |  | 2.74 | High Complexity
|  | Gold Standard Diagnostics | B. burgdorferiVlsE-OspC IgG/IgM |  | Serum (119364003^Serum specimen^SCT) | Cutoff = OD Calibrator*Correction Factor from the calibrator label Units (qualitative) = (OD Controls of samples / Cutoff)*10 <9.0 units = negative (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) 9.0 - 11 units = equivocal (42425007^Equivocal^SCT) >11 units = positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 101357-2 | Borrelia burgdorferi.VlsE+OspC IgG+IgM Ab [Presence] in Serum by Immunoassay | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel |  |  |  | Borrelia burgdorferi VLSE-OspC IgG/IgM ELISA Test Kit_Gold Standards Diagnostics, INC | MNT | No specific equipment | N/A | Borrelia burgdorferi VlsE and OspC Ab.IgG+IgM | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | Gold Standard Diagnostics | B. burgdorferi IgG/IgM |  | Serum (119364003^Serum specimen^SCT) | Cutoff = OD Calibrator*Correction Factor from the calibrator label Units (qualitative) = (OD Controls of samples / Cutoff)*10 <9.0 units = negative (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) 9.0 - 11 units = equivocal (42425007^Equivocal^SCT) >11 units = positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 83081-0 | Borrelia burgdorferi IgG+IgM Ab [Presence] in Serum by Immunoassay | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel |  |  |  | Borrelia burgdorferi IgG/IgM ELISA Test Kit_Gold Standards Diagnostics, INC | MNT | No specific equipment | N/A | Borrelia burgdorferi Ab.IgG+IgM | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | Gold Standard Diagnostics | B. burgdorferi IgG |  | Serum (119364003^Serum specimen^SCT) | Cutoff = OD Calibrator*Correction Factor from the calibrator label Units (qualitative) = (OD Controls of samples / Cutoff)*10 <9.0 units = negative (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) 9.0 - 11 units = equivocal (42425007^Equivocal^SCT) >11 units = positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel |  |  |  | Borrelia burgdorferi IgG ELISA Test Kit_Gold Standards Diagnostics, INC | MNT | No specific equipment | N/A | Borrelia burgdorferi Ab.IgG | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | Gold Standard Diagnostics | B. burgdorferi IgM |  | Serum (119364003^Serum specimen^SCT) | Cutoff = OD Calibrator*Correction Factor from the calibrator label Units (qualitative) = (OD Controls of samples / Cutoff)*10 <9.0 units = negative (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) 9.0 - 11 units = equivocal (42425007^Equivocal^SCT) >11 units = positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel |  |  |  | Borrelia burgdorferi IgM ELISA Test Kit_Gold Standards Diagnostics, INC | MNT | No specific equipment | N/A | Borrelia burgdorferi Ab.IgM | PrThr | Pt | Ser | Ord | IA |  | 2.74 | High Complexity
|  | Gold Standard Diagnostics | B. burgdorferi Exposure Summary |  | Serum (119364003^Serum specimen^SCT) | IgG- and IgM - (260415000^Not detected^SCT – Clinical interpretation: “No evidence of infection.” - this should be sent as NTE after this OBX) IgG – and IgM + (711621000124102^Borrelia burgdorferi immunoglobulin M antibody detected^SCT - the clinical interpretation is: "Probable infection at an early stage." - this should be sent as NTE after this OBX) IgG+ and IgM- (711631000124104^Borrelia burgdorferi immunoglobulin G antibody detected^SCT) - the clinical interpretation is: "Probable infection at any stage." - this should be sent as NTE after this OBX IgG+ and IgM+ (711611000124105^Borrelia burgdorferi immunoglobulin G and M antibodies detected^SCT - the clinical interpretation is: "Probable acute infection." - this should be sent as NTE after this OBX) | 101358-0 | Borrelia burgdorferi Ab [Interpretation] in Serum by IA.MTTT | 101352-3 | Borrelia burgdorferi Ab.IgG & IgM with reflex to immunoassay panel | [7 empty] | Borrelia burgdorferi Ab.IgG & IgM | Imp | Pt | Ser | Nom | IA |  | 2.74 | High Complexity
|  | Trinity Biotech | MarDx Borrelia burgdorferi EIA IgG + IgM | IgG + IgM Ab | Serum (119364003^Serum specimen^SCT) | Lyme Index Value: >=1.2 (Positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT). "This result should be supplemented by a second tier test that is more specific for antibodies to B. burgdorferi and should not be reported until second step testing is complete." Lyme Index Value: <1.0 (Negative (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT)) Lyme Index Value: >=1.0 and <1.2 (Equivocal (42425007^Equivocal^SCT) "Equivocal result is presumptive evidence of possible exposure to B. burgdorferi. The result should be supplemented by a second-tier test that is more specific for antibodies to B. burgdorferi and should not be reported until second-step testing is complete." | 83081-0 | Borrelia burgdorferi IgG+IgM Ab [Presence] in Serum by Immunoassay | 83081-0 | Borrelia burgdorferi IgG+IgM Ab [Presence] in Serum by Immunoassay | EIA for the detection of IgG and IgM antibody to Borrelia burgdorferi 96 tests |  |  | MarDx Borrelia burgdorferi EIA IgG + IgM_Trinity Biotech | MNT | <no specific equipment - Use of a single (450nm) wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgG+IgM | PrThr | Pt | Ser | Ord | IA |  | 2.76 | High Complexity | EIA | Indirect Enzyme Immunoassay
|  | Trinity Biotech | MarDx Borrelia burgdorferi EIA IgG | IgG Ab | Serum (119364003^Serum specimen^SCT) | Lyme Index Value: >=1.2 (Positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT). "This result should be supplemented by a second tier test that is more specific for antibodies to B. burgdorferi and should not be reported until second step testing is complete." Lyme Index Value: <0.8 (Negative (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT)) Lyme Index Value: >=0.8 and <1.2 (equivocal (42425007^Equivocal^SCT) "Equivocal result is presumptive evidence of possible exposure to B. burgdorferi. The result should be supplemented by a second-tier test that is more specific for antibodies to B. burgdorferi and should not be reported until second-step testing is complete." | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay | EIA for the detection of IgG antibody to Borrelia burgdorferi 96 tests |  |  | MarDx Borrelia burgdorferi EIA IgG_Trinity Biotech | MNT | <no specific equipment - Use of a single (450nm) wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgG | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | EIA | Indirect Enzyme Immunoassay
|  | Trinity Biotech | MarDx Borrelia burgdorferi EIA IgM | IgM Ab | Serum (119364003^Serum specimen^SCT) | Lyme Index Value: >=1.2 (Positive (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT). "This result should be supplemented by a second tier test that is more specific for antibodies to B. burgdorferi and should not be reported until second step testing is complete." Lyme Index Value: <0.8 (Negative (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT)) Lyme Index Value: >=0.8 and <1.2 (equivocal (42425007^Equivocal^SCT) "Equivocal result is presumptive evidence of possible exposure to B. burgdorferi. The result should be supplemented by a second-tier test that is more specific for antibodies to B. burgdorferi and should not be reported until second-step testing is complete." | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay | EIA for the detection of IgG antibody to Borrelia burgdorferi 480 tests |  |  | MarDx Borrelia burgdorferi EIA IgM_Trinity Biotech | MNT | <no specific equipment - Use of a single (450nm) wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgM | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | EIA | Indirect Enzyme Immunoassay
|  | DiaSorin | LIAISON Lyme IgM | Lyme IgM Result | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Index = < 0.90 = Negative = Absence of detectable Borrelia burgdorferi IgM antibodies. A negative result does not exclude the possibility of Borrelia burgdorferi infection. If exposure to B. burgdorferi is suspected, a second sample should be collected and tested two to four weeks later. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Index = ≥ 0.90 and < 1.10 = Equivocal = Testing guidelines recommend that all equivocal samples be tested further following the standard two-tier test or modified two-tier test methodologies. (42425007^Equivocal^SCT) Index = ≥ 1.10 = Positive = Presumptive detection of anti-Borrelia burgdorferi IgM antibodies. Current testing guidelines recommend that all positive results be supplemented further following the standard two-tier test or modified two-tier test methodologies. (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme IgM_DiaSorin Inc | MNT | LIAISON XL Analyzer_DiaSorin | MNI | Borrelia burgdorferi Ab.IgM | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | CLIA | Indirect chemiluminescence Immunoassay
|  | DiaSorin | LIAISON Lyme IgM | Lyme IgM Result | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Index = < 0.90 = Negative = Absence of detectable Borrelia burgdorferi IgM antibodies. A negative result does not exclude the possibility of Borrelia burgdorferi infection. If exposure to B. burgdorferi is suspected, a second sample should be collected and tested two to four weeks later. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Index = ≥ 0.90 and < 1.10 = Equivocal = Testing guidelines recommend that all equivocal samples be tested further following the standard two-tier test or modified two-tier test methodologies. (42425007^Equivocal^SCT) Index = ≥ 1.10 = Positive = Presumptive detection of anti-Borrelia burgdorferi IgM antibodies. Current testing guidelines recommend that all positive results be supplemented further following the standard two-tier test or modified two-tier test methodologies. (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme IgM_DiaSorin Inc | MNT | LIAISON XS Analyzer_DiaSorin | MNI | Borrelia burgdorferi Ab.IgM | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | CLIA | Indirect chemiluminescence Immunoassay
|  | DiaSorin | LIAISON Lyme IgG | Lyme IgG Result | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Index = < 0.90 = Negative = Absence of detectable Borrelia burgdorferi IgM antibodies. A negative result does not exclude the possibility of Borrelia burgdorferi infection. If exposure to B. burgdorferi is suspected, a second sample should be collected and tested two to four weeks later. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Index = ≥ 0.90 and < 1.10 = Equivocal = Testing guidelines recommend that all equivocal samples be tested further following the standard two-tier test or modified two-tier test methodologies. (42425007^Equivocal^SCT) Index = ≥ 1.10 = Positive = Presumptive detection of anti-Borrelia burgdorferi IgM antibodies. Current testing guidelines recommend that all positive results be supplemented further following the standard two-tier test or modified two-tier test methodologies. (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme IgG_DiaSorin Inc | MNT | LIAISON XL Analyzer_DiaSorin | MNI | Borrelia burgdorferi Ab.IgG | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | CLIA | Indirect chemiluminescence Immunoassay
|  | DiaSorin | LIAISON Lyme IgG | Lyme IgG Result | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Index = < 0.90 = Negative = Absence of detectable Borrelia burgdorferi IgM antibodies. A negative result does not exclude the possibility of Borrelia burgdorferi infection. If exposure to B. burgdorferi is suspected, a second sample should be collected and tested two to four weeks later. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Index = ≥ 0.90 and < 1.10 = Equivocal = Testing guidelines recommend that all equivocal samples be tested further following the standard two-tier test or modified two-tier test methodologies. (42425007^Equivocal^SCT) Index = ≥ 1.10 = Positive = Presumptive detection of anti-Borrelia burgdorferi IgM antibodies. Current testing guidelines recommend that all positive results be supplemented further following the standard two-tier test or modified two-tier test methodologies. (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme IgG_DiaSorin Inc | MNT | LIAISON XS Analyzer_DiaSorin | MNI | Borrelia burgdorferi Ab.IgG | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | CLIA | Indirect chemiluminescence Immunoassay
|  | ViraMed Biotech AG | Borrelia B31 IgM ViraStripe | Borrelia B31 IgM | Serum (119364003^Serum specimen^SCT) | Positive (at least two significant (+) bands (kD) from 41, 39, 23)(10828004^Positive^SCT - if allowable should map to the recommended term, which is 260373001^Detected^SCT). IgM-antibodies against Borrelia species detectable. Presumptive evidence of B. burgdorferi infection. Negative (No bands (kD) or less than two significant bands) (260385009^Negative^SCT - if allowable should map to the recommended term, which is 260415000^Not detected^SCT)). No IgM-antibodies against Borrelia species detectable. In case of a clinically based suspicion of an infection with Borrelia: check additionally for IgG-antibodies and possibly check a second sample for IgG antibodies and IgM-antibodies after 2-3 weeks. | 104602-8 | Borrelia burgdorferi B31 IgM Ab [Presence] in Serum by Line blot | 104602-8 | Borrelia burgdorferi B31 IgM Ab [Presence] in Serum by Line blot |  |  | V-BBSMUS/V-BBSMDS | ViraMed Biotech AG_Borrelia B31 IgM ViraStripe | MNT | <no specific equipment - Use of a line blot membrane strip well/western line blot assay processor.> | N/A | Borrelia burgdorferi B31 Ab.IgM | PrThr | Pt | Ser | Ord | Line blot |  | pre-2.78 | High Complexity | Line blot assay | Line blot assay
|  | ViraMed Biotech AG | Borrelia B31 IgG ViraStripe | Borrelia B31 IgG | Serum (119364003^Serum specimen^SCT) | Positive (at least five significant (+) bands (kD) from 93, 66, 58, 45, 41, 39, 30, 28, 23, 18)(10828004^Positive^SCT - if allowable should map to the recommended term, which is 260373001^Detected^SCT) IgG-antibodies against Borrelia species detectable. Presumptive evidence of B. burgdorferi infection. Negative (No bands (kD) or less than five significant bands) (260385009^Negative^SCT - if allowable should map to the recommended term, which is 260415000^Not detected^SCT)). No IgG-antibodies against Borrelia species detectable. In case of a clinically based suspicion of an infection with Borrelia: check additionally for IgM-antibodies and possibly check a second sample for IgG antibodies and IgM-antibodies after 2-3 weeks. | 104603-6 | Borrelia burgdorferi B31 IgG Ab [Presence] in Serum by Line blot | 104603-6 | Borrelia burgdorferi B31 IgG Ab [Presence] in Serum by Line blot |  |  | V-BBSGUS/V-BBSGDS | ViraMed Biotech AG_Borrelia B31 IgG ViraStripe | MNT | <no specific equipment - Use of a line blot membrane strip well/western line blot assay processor.> | N/A | Borrelia burgdorferi B31 Ab.IgG | PrThr | Pt | Ser | Ord | Line blot |  | pre-2.78 | High Complexity | Line blot assay | Line blot assay
|  | BIOMÉRIEUX | VIDAS Lyme IgG II (LYG) | IgG | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Cutoff is based on the receiver operating characteristic (ROC) curve analysis. Cut-off value for VIDAS Lyme IgG II was determined to be be equal to 0.20. Result index <0.20 = Negative (260385009^Negative^SCT - if allowable should map to the recommended term, which is 260415000^Not detected^SCT). No detectable anti-B. burgdorferi IgG antibodies detected. Result index >=0.20 = Positive (10828004^Positive^SCT - if allowable should map to the recommended term, which is 260373001^Detected^SCT). Presumptive detection of anti-B. burgdorferi IgG antibodies. The specimen should be tested further by an IgG Western Blot for supplementary information. | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay | This test is intended for qualitative detection only. |  |  | VIDAS Lyme IgG II (LYG)_BIOMÉRIEUX | MNT | <no specific equipment - Use of a single (450nm) wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgG | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | Enzyme Immunoassay | Enzyme Immunoassay
|  | BIOMÉRIEUX | VIDAS Lyme IgM II (LYM) | IgM | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Cutoff is based on the receiver operating characteristic (ROC) curve analysis. Cut-off value for VIDAS Lyme IgM II was determined to be be equal to 0.12. Result index <0.12 = Negative (260385009^Negative^SCT - if allowable should map to the recommended term, which is 260415000^Not detected^SCT). No detectable anti-B. burgdorferi IgM antibodies detected. Result index >=0.12 and <0.32 = Equivocal (42425007^Equivocal^SCT). Imprecise detection of anti-B. burgorferi IgM antibodies. The specimen should be tested further by an IgM Western Blot for supplementary information. Result index >=0.32 = Positive (10828004^Positive^SCT - if allowable should map to the recommended term, which is 260373001^Detected^SCT). Presumptive detection of anti-B. burgdorferi IgM antibodies. The specimen should be tested further by an IgM Western Blot for supplementary information. | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay | This test is intended for qualitative detection only. |  |  | VIDAS Lyme IgM II (LYM)_BIOMÉRIEUX | MNT | <no specific equipment - Use of a single (450nm) wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgM | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | Enzyme Immunoassay | Enzyme Immunoassay
|  | Zeus Scientific, Inc | B. burgdorferi IgG/IgM Test System | IgG and IgM antibodies | Serum (119364003^Serum specimen^SCT) | Index Value/OD Ratio Negative Specimens ≤0.90 (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT). No significant amount of IgM and/or IgG antibodies to B. burgdorferi detected. An additional sample should be tested within four to six weeks if early infection is suspected. Equivocal Specimens 0.91 to 1.09 (42425007^Equivocal^SCT) Positive Specimens ≥1.10 (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT). Presumptively positive for IgG/IgM antibody to B. burgdorferi. Per current recommendations, the result cannot be further interpreted without supplemental Western Blot testing. Results should not be reported until the supplemental testing is completed. | 83081-0 | Borrelia burgdorferi IgG+IgM Ab [Presence] in Serum by Immunoassay | 83081-0 | Borrelia burgdorferi IgG+IgM Ab [Presence] in Serum by Immunoassay |  |  |  | B. burgdorferi IgG/IgM Test System_Zeus Scientific, Inc | MNT | <no specific equipment - Use of a single (450nm), or dual (450/620 - 650nm), wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgG+IgM | PrThr | Pt | Ser | Ord | IA |  | 2.76 | High Complexity | ELISA | ELISA
|  | Zeus Scientific, Inc | B. burgdorferi IgG Test System | IgG antibodies | Serum (119364003^Serum specimen^SCT) | Index Value/OD Ratio Negative Specimens ≤0.90 (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT). If exposure to to B. burgdorferi is suspected, a second sample should be collected and tested two to four weeks later. Equivocal Specimens 0.91 to 1.09 (42425007^Equivocal^SCT). Evaluate repeatedly equivocal specimen by the second step IgG and/or IgM Western blot. Positive Specimens ≥1.10 (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT). Indicates that antibodies specific to B burgdorferi were detected. This indicates presumptive evidence of probable exposure. Per current recommendations, the result cannot be further interpreted without supplemental Western Blot testing. Results should not be reported until the supplemental testing is completed. | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay | 16480-6 | Borrelia burgdorferi IgG Ab [Presence] in Serum or Plasma by Immunoassay |  |  |  | B. burgdorferi IgG Test System_Zeus Scientific, Inc | MNT | <no specific equipment - Use of a single (450nm), or dual (450/620 - 650nm), wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgG | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | ELISA | ELISA
|  | Zeus Scientific, Inc | B. burgdorferi IgM Test System | IgM antibodies | Serum (119364003^Serum specimen^SCT) | Index Value/OD Ratio Negative Specimens ≤0.90 (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT). No detectable IgM antibody; result does not exclude B. burgdorferi infection. An additional sample should be tested within four to six weeks if early infection is suspected. Equivocal Specimens 0.91 to 1.09 (42425007^Equivocal^SCT). Current recommendations state that equivocal results should be followed by supplemental Western Blot testing. Results should not be reported until the supplemental testing is completed. Positive Specimens ≥1.10 (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT). IgM antibody to B. burgdorferi presumptively detected. Per current recommendations, the result cannot be further interpreted without supplemental Western Blot testing. Do not report results until supplemental testing is completed. | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay | 40612-4 | Borrelia burgdorferi IgM Ab [Presence] in Serum or Plasma by Immunoassay |  |  |  | B. burgdorferi IgM Test System_Zeus Scientific, Inc | MNT | <no specific equipment - Use of a single (450nm), or dual (450/620 - 650nm), wavelength reader is acceptable.> | N/A | Borrelia burgdorferi Ab.IgM | PrThr | Pt | Ser/Plas | Ord | IA |  | 2.76 | High Complexity | ELISA | ELISA

## LOINC Mapping Columns
LIVD Spreadsheet Guidance
· Manufacturer – the name of the Instrument manufacturer or Testkit manufacturer
· Model – the name of the Instrument or the Testkit
· Vendor Analyte Name – the human-readable text the vendor uses to identify the analyte. The text might be displayed by the instrument or could be used within an assay insert or instructions for use.
· Vendor Specimen Description – the human-readable text that provides information about the specimen used for the test, such as “Serum or Plasma.” The field documents the vendor description of the specimen used for the IVD test. The field may contain multiple specimen types. For this catalog, the SNOMED-CT mapping for each Specimen is provided in parenthesis
o Serum (119364003^Serum specimen^SCT)
o Plasma (119361006^Plasma specimen^SCT)
· Vendor Result Description – the human-readable text that provides information about the result that is produced.
o For non-numeric results, this field contains the possible result values, along with the SNOMED-CT mapping
§ Positive (10828004^Positive^SCT)
§ Negative (260385009^Negative^SCT)
o For numeric results and associated units of measure this field describes the result by including a representative unit of measure, preferably represented as a UCUM unit.
· Test Performed LOINC® Code – the appropriate LOINC code for the Test Performed mapping
· Test Performed LOINC Long Name – the long name for the Test Performed LOINC Code
· Test Ordered LOINC Code – the appropriate LOINC code for the Test Ordered mapping
· Test Ordered LOINC Long Name– the long name for the Test Ordered LOINC Order
· Vendor Comment – human-readable text for clarification, such as “This is a STAT (prioritized) version of the test”
· Testkit Name ID – the Testkit Device Identifier if known, otherwise the concatenation of the <Diagnostic (Letter of Authorization)>_<Manufacturer> fields from the FDA EUA Website
· Testkit Name ID Type – GTIN when the Device Identifier is known, otherwise EUA when the Emergency Use Authorization identifiers are used
· Vendor Analyte Code – one of two possible values
o For an automated test, it contains Vendor Transmission Code used by the instrument when sending the test result to a health information system, such as an LIS.
o For a manual test, it is the Vendor Analyte Identifier for the test result produced by the Test Kit.
· Vendor Reference ID – a vendor identifier, such as an identifier that can be used to locate the associated assay insert published by the vendor. This attribute may contain the material number used to order the product from the manufacturer.
· Equipment UID – when available, the Device Identifier of the instrument. If not known, then the concatenation of the Model and Manufacturer, <Model>_<Manufacturer>
· Equipment UID Type – GTIN when the Device Identifier is provided, otherwise MNI when model and manufacturer are provided
· Component – the LOINC Component/analyte
· Property – the LOINC Kind of Property
· Time – the LOINC Time Aspect
· System – The LOINC System (Sample)
· Scale – the LOINC Type of Scale
· Method – the LOINC Type of Method
· Publication Version ID – human-readable information used to differentiate mapping publication versions
· LOINC Version ID – the version of LOINC used to establish the LOINC mapping

## Instruments
Equipment ID | Equipment Name | Equipment Name Type | Equipment UID | Equipment UID Type | Equipment Comment | Equipment Sort
1000 | (Test kit device) | (NA) |  |  | This test is run on the rapid test kit device, not a separate instrument. | 1005
1231 | a luminometer_NA | MNI |  |  | Some Package Inserts just indicate that the test is run on "a luminometer" | 2001
1230 | a microplate reader_NA | MNI |  |  | Some Package Inserts just indicate that the test is run on "a microplate reader" | 2004
1241 | an ELISA spectrophotometer_NA | MNI |  |  | Some Package Inserts just indicate that the test is run on "an ELISA reader" or "an ELISA spectrophotometer" | 2007
1179 | ABI 3500xL Dx Genetic Analyzer_Thermo Fisher | MNI |  |  |  | 2010
1066 | ABI 3500xL Genetic Analyzer_Thermo Fisher | MNI |  |  |  | 2020
1065 | ABI 3730xl DNA Analyzer_Thermo Fisher | MNI |  |  |  | 2030
1002 | ABI 7500 FAST Dx Real-Time PCR_Thermo Fisher | MNI | 10190302000222 | DII |  | 2040
1074 | ABI 7500 FAST Real-Time PCR_Thermo Fisher | MNI |  |  |  | 2050
1075 | ABI 7500 Real-Time PCR_Thermo Fisher | MNI |  |  |  | 2060
1076 | ABI 7900HT Fast Real-Time PCR_Thermo Fisher | MNI |  |  |  | 2070
1064 | ABI GeneAmp PCR System 9700_Thermo Fisher | MNI |  |  |  | 2080
1077 | ABI MiniAmp Thermal Cycler_Thermo Fisher | MNI |  |  |  | 2090
1063 | ABI PRISM 7000 Detection System_Thermo Fisher | MNI |  |  |  | 2100
1067 | ABI QuantStudio 12K Flex_Thermo Fisher | MNI |  |  |  | 2110
1062 | ABI QuantStudio 3 Real-Time PCR_Thermo Fisher | MNI |  |  |  | 2120
1061 | ABI QuantStudio 3D Digital PCR_Thermo Fisher | MNI |  |  |  | 2130
1060 | ABI QuantStudio 5 Dx_Thermo Fisher | MNI |  |  |  | 2140
1001 | ABI QuantStudio 5 Real-Time PCR_Thermo Fisher | MNI | 10190302015592 | DII |  | 2150
1078 | ABI QuantStudio 6 Flex_Thermo Fisher | MNI |  |  |  | 2160
1128 | ABI QuantStudio 6 Pro_Thermo Fisher | MNI |  |  |  | 2170
1079 | ABI QuantStudio 7 Flex_Thermo Fisher | MNI |  |  |  | 2180
1080 | ABI QuantStudio 7 Pro_Thermo Fisher | MNI |  |  |  | 2190
1081 | ABI QuantStudio Dx_Thermo Fisher | MNI |  |  |  | 2200
1132 | ABI StepOne Real-Time PCR_Thermo Fisher | MNI |  |  |  | 2210
1133 | ABI StepOnePlus Real-Time PCR_Thermo Fisher | MNI |  |  |  | 2220
1068 | ABI Veriti Dx Thermal Cycler_Thermo Fisher | MNI |  |  |  | 2230
1082 | ABI Veriti Thermal Cycler_Thermo Fisher | MNI |  |  |  | 2240
1083 | ABI ViiA 7 PCR System_Thermo Fisher | MNI |  |  |  | 2250
1139 | Access 2 Immunoassay System_BeckmanCoulter | MNI |  |  |  | 2260
1138 | Access Immunoassay System_BeckmanCoulter | MNI |  |  |  | 2270
1161 | Accula Dock_Mesa Biotech | MNI |  |  |  | 2280
1178 | ADVIA Centaur CP system_Siemens | MNI |  |  |  | 2290
1069 | ADVIA Centaur XP system_Siemens | MNI |  |  |  | 2300
1070 | ADVIA Centaur XPT system_Siemens | MNI |  |  |  | 2310
1059 | Alinity i_Abbott | MNI | 00380740137366 | DII |  | 2320
1058 | Alinity m_Abbott | MNI | 00884999048034 | DII |  | 2330
1264 | Araya System_Biosearch | MNI |  |  |  | 2335
1057 | Architect i1000SR_Abbott | MNI | 00380740001537 | DII |  | 2340
1056 | Architect i2000SR_Abbott | MNI | 00380740006235 | DII |  | 2350
1071 | Aria Mx_Agilent Technologies | MNI |  |  |  | 2360
1084 | Atellica IM Analyzer_Siemens | MNI |  |  |  | 2370
1085 | BD MAX System_Becton Dickinson | MNI | 00382904419165 | DII |  | 2380
1159 | BD Veritor Plus System_Becton Dickinson | MNI |  |  |  | 2390
1087 | BioCode MDx-3000_Applied BioCode | MNI |  |  |  | 2400
1172 | BIO-FLASH Instrument_Inova | MNI |  |  |  | 2410
1088 | Biomark HD system_Standard BioTools Inc. | MNI |  |  |  | 2420
1089 | Biomeme Franklin three9 Thermocycler_Biomeme | MNI |  |  |  | 2430
1180 | BioPlex 2200 Instrument_BioRad | MNI |  |  |  | 2440
1055 | BioTek 800 TS Absorbance Reader_BioTek | MNI |  |  |  | 2450
1252 | Bravo BenchCel Workstation_Agilent | MNI |  |  |  | 2460
1157 | C1000 Dx Thermal Cycler_BioRad | MNI |  |  |  | 2470
1091 | CFX Connect Real-Time PCR System_BioRad | MNI |  |  |  | 2480
1154 | CFX Opus 384_BioRad | MNI |  |  |  | 2490
1153 | CFX Opus 96_BioRad | MNI |  |  |  | 2500
1097 | CFX384 Real-Time PCR_BioRad | MNI |  |  |  | 2510
1131 | CFX384 Touch Real-Time PCR_BioRad | MNI |  |  |  | 2520
1156 | CFX96 DX Real-Time PCR System ORM_BioRad | MNI |  |  | ORM is Optical Reading Module | 2530
1092 | CFX96 DX Real-Time PCR System_BioRad | MNI |  |  |  | 2540
1054 | CFX96 Real-Time PCR System_BioRad | MNI |  |  |  | 2550
1053 | CFX96 Touch Deep Well Real-Time PCR_BioRad | MNI |  |  |  | 2560
1093 | CFX96 Touch Real-Time PCR System_BioRad | MNI |  |  |  | 2570
1130 | Clear Dx system_Clear Labs | MNI |  |  |  | 2580
1148 | Clip Analyzer_Luminostics | MNI |  |  |  | 2590
1094 | cobas 6800 System_Roche | MNI |  |  |  | 2600
1095 | cobas 8800 System_Roche | MNI |  |  |  | 2610
1147 | cobas e 402 analyzer_Roche | MNI | 7613336165300 | DII | Roche calls the Equipment UID type 'GTIN' | 2620
1146 | cobas e 411 analyzer_Roche | MNI | 4015630924721 | DII | Roche calls the Equipment UID type 'GTIN' | 2640
1051 | cobas e 601 analyzer_Roche | MNI | 4015630924714 | DII | Roche calls the Equipment UID type 'GTIN' | 2650
1050 | cobas e 602 analyzer_Roche | MNI | 4015630929764 | DII | Roche calls the Equipment UID type 'GTIN' | 2660
1049 | cobas e 801 analyzer_Roche | MNI | 4015630946198 | DII | Roche calls the Equipment UID type 'GTIN' | 2670
1003 | cobas Liat Analyzer_Roche | MNI |  |  |  | 2680
1096 | cobas z 480 Instrument_Roche | MNI |  |  |  | 2690
1048 | cobas 8000 modular analyzer_Roche | MNI |  |  |  | 2700
1047 | CoDx Box_Co-Diagnostics | MNI |  |  |  | 2710
1046 | Cue Health Monitoring System_Cue Health | MNI |  |  |  | 2720
1255 | DASH Analyzer_Minute Molecular Diagnostics | MNI |  |  |  | 2730
1045 | Dimension EXL 200_Siemens | MNI |  |  |  | 2740
1044 | Dimension EXL with LOCI Module_Siemens | MNI |  |  |  | 2750
1043 | Dimension Vista System_Siemens | MNI |  |  |  | 2760
1042 | DropX-2000 digital PCR system_PreciGenome | MNI |  |  |  | 2770
1099 | DTPrime5_DNA Technologie | MNI |  |  |  | 2780
1265 | DxHub Instrument_DxLab | MNI |  |  |  | 2785
1141 | Dynex Agility System_Dynex Technologies | MNI |  |  |  | 2790
1177 | Dynex DS2 System_Dynex Technologies | MNI |  |  |  | 2800
1176 | Dynex DSX System_Dynex Technologies | MNI |  |  |  | 2810
1100 | DZ-lite 3000 Plus Analyzer_Diazyme | MNI |  |  |  | 2820
1257 | Eco 48 Real-Time qPCR System_PCRmax | MNI |  |  |  | 2830
1101 | EdvoCycler 2_Edvotek | MNI |  |  |  | 2840
1102 | EdvoCycler Jr_Edvotek | MNI |  |  |  | 2850
1041 | Encompass MDx Workstation_Rheonix | MNI |  |  |  | 2860
1237 | EUROLabWorkstation ELISA_EUROIMMUN | MNI |  |  |  | 2870
1040 | FilmArray 2.0 System_BioFire | MNI |  |  |  | 2880
1039 | FilmArray Torch System_BioFire | MNI |  |  |  | 2890
1104 | FLEXMAP 3D Sys(MFIdivider on)_Luminex Corp | MNI |  |  |  | 2900
1103 | FLEXMAP 3D System_Luminex Corp | MNI |  |  |  | 2910
1038 | FLEXMAP dual laser cytometer_Luminex Corp | MNI |  |  |  | 2920
1037 | Franklin Real-Time PCR Thermocycler_Biomeme | MNI |  |  |  | 2930
1136 | FREND System_NanoEntek | MNI |  |  |  | 2940
1036 | GeneXpert Dx System_Cepheid | MNI |  |  |  | 2950
1035 | GeneXpert Infinity System_Cepheid | MNI |  |  |  | 2960
1145 | GeneXpert Omni System_Cepheid | MNI |  |  |  | 2970
1034 | GeneXpert Xpress System_Cepheid | MNI |  |  |  | 2980
1105 | Genie HT instrument_OptiGene | MNI |  |  |  | 2990
1033 | GenMark ePlex instrument_GenMarkDx | MNI |  |  |  | 3000
1129 | GridION nanopore sequencer_Oxford Nanopore | MNI |  |  |  | 3010
1171 | Hydrocycler2_Biosearch | MNI |  |  |  | 3020
1032 | ID NOW Instrument_Abbott | MNI | 00811877010616 | DII |  | 3030
1106 | Illumina NovaSeq 6000 SP_Illumina | MNI |  |  |  | 3040
1170 | IntelliQube qPCR system_Biosearch | MNI |  |  |  | 3050
1158 | iSYS Multi-Discipline Automated System_IDS | MNI |  |  | IDS is Immunodiagnostic Systems | 3060
1030 | LIAISON MDX_DiaSorin | MNI |  |  |  | 3070
1135 | LIAISON XL Analyzer_DiaSorin | MNI |  |  |  | 3080
1234 | Liberty16 mobile real time PCR_Ubiquitome | MNI |  |  |  | 3090
1108 | LightCycler 480 Instrument II_Roche | MNI |  |  |  | 3100
1107 | LightCycler 480 Instrument_Roche | MNI |  |  |  | 3110
1029 | LightCycler 96 Instrument_Roche | MNI |  |  |  | 3120
1109 | Luminex 200 System_Luminex Corp | MNI |  |  |  | 3130
1028 | Luminex ARIES M1 System_Luminex Corp | MNI |  |  |  | 3140
1027 | Luminex ARIES System_Luminex Corp | MNI |  |  |  | 3150
1112 | Luminex MAGPIX System_Luminex Corp | MNI |  |  |  | 3160
1110 | LumiraDx Platform_LumiraDx | MNI |  |  |  | 3170
1026 | m2000 RealTime System_Abbott | MNI | 00884999031814 | DII |  | 3180
1111 | MAGLUMI 2000 series analyzer_Shenzhen | MNI |  |  |  | 3190
1113 | MassARRAY System_Agena Bioscience | MNI |  |  |  | 3200
1140 | Maverick Diagnostic System_Genalyte | MNI |  |  |  | 3210
1098 | MD-Bio Heater_Detectachem | MNI |  |  |  | 3220
1025 | Mic PCR cycler_Bio Molecular Systems | MNI |  |  |  | 3230
1114 | MINI VIDAS_bioMérieux | MNI |  |  |  | 3240
1142 | MiSeq System_Illumina | MNI |  |  |  | 3250
1127 | MosaIQ 125_Quotient | MNI |  |  |  | 3260
1115 | MS Fast Analyzer_Sophonix | MNI |  |  |  | 3270
1024 | NeuMoDx 288 Molecular System_NeuMoDx | MNI |  |  |  | 3280
1023 | NeuMoDx 96 Molecular System_NeuMoDx | MNI |  |  |  | 3290
1022 | NextSeq 500 Sequencing System_Illumina | MNI |  |  |  | 3300
1021 | NextSeq 550 Sequencing System_Illumina | MNI |  |  |  | 3310
1144 | NextSeq 550Dx Sequencing System_Illumina | MNI |  |  |  | 3320
1031 | NovaSeq 6000 Sequencing System_Illumina | MNI |  |  |  | 3330
1116 | Panther Fusion System_Hologic | MNI |  |  |  | 3340
1117 | Panther System_Hologic | MNI |  |  |  | 3350
1152 | Phadia 250 instrument_Phadia AB | MNI |  |  |  | 3360
1020 | Platelia system_BioRad | MNI |  |  |  | 3370
1260 | PNY-1000 instrument_InspectIR | MNI |  |  |  | 3380
1019 | PowerGene 9600 Plus Real-time PCR_Atila | MNI |  |  |  | 3390
1155 | Procleix Panther System_Grifols | MNI |  |  |  | 3400
1229 | QIAreach eHub_QIAGEN GmbH | MNI |  |  |  | 3410
1018 | QIAstat-Dx Analyzer 1.0_QIAGEN GmbH | MNI |  |  |  | 3420
1169 | Qorvo Omnia Instrument_Qorvo | MNI |  |  |  | 3430
1073 | qTower3 G_Analytik Jena GmbH | MNI |  |  |  | 3440
1174 | qTower3_Analytik Jena GmbH | MNI |  |  |  | 3450
1267 | qTower3 84 G_Analytik Jena GmbH | MNI |  |  |  | 3452
1266 | qTower3 84_Analytik Jena GmbH | MNI |  |  |  | 3453
1227 | Q-View Imager LS_Quansys | MNI |  |  |  | 3460
1017 | Q-View Imager Pro_Quansys | MNI |  |  |  | 3470
1016 | QX200 Droplet Digital PCR System_BioRad | MNI |  |  |  | 3480
1119 | Rapid COVID-19 IgM/IgG Combo Kit_Megna Health | MNI |  |  |  | 3490
1015 | Real Time Digital PCR Instrument_Gnomegen | MNI |  |  |  | 3500
1245 | Revogene instrument_Meridian | MNI |  |  |  | 3510
1120 | Rotor-Gene Q 5plex HRM_QIAGEN GmbH | MNI |  |  |  | 3520
1014 | Rotor-Gene Q 5plex_QIAGEN GmbH | MNI |  |  |  | 3530
1134 | Rotor-Gene Q MDx_QIAGEN GmbH | MNI |  |  |  | 3540
1118 | Rotor-Gene Q_QIAGEN GmbH | MNI |  |  |  | 3550
1261 | Sal6830 Point of Care PCR System_MicroGEM | MNI |  |  |  | 3560
1143 | Sampinute Analyzer_Celltrion | MNI |  |  |  | 3570
1173 | SensoSpot Microarray Analyzer_Sensovation AG | MNI |  |  |  | 3580
1137 | SensoSpot Microarray Scanner_Sensovation | MNI |  |  |  | 3590
1121 | Sentosa SA201 Real-Time PCR_Vela Diagnostics | MNI |  |  |  | 3600
1160 | Silaris Dock_Sekisui Diagnostics | MNI |  |  |  | 3610
1150 | Simoa HD-X Analyzer_Quanterix | MNI |  |  |  | 3620
1013 | SLAN-96P_Shanghai Hongshi | MNI |  |  |  | 3630
1012 | Sofia 2 Instrument_Quidel | MNI |  |  |  | 3640
1011 | Sofia Instrument_Quidel | MNI |  |  |  | 3650
1149 | Solana Instrument_Quidel | MNI |  |  | Solana is a bench top instrument that combines Quidel's proprietary helicase-dependent amplification (HDA) with fluorescence detection | 3660
1151 | Solas 8_MatMaCorp | MNI |  |  | The Solas 8 is a portable, four channel fluorescence detection system | 3670
1238 | Sprinter XL_EUROIMMUN | MNI |  |  |  | 3680
1072 | Stratagene Mx3005P_Agilent Technologies | MNI |  |  |  | 3690
1010 | Synergy NEO2 Hybrid Multi-Mode Reader_BioTek | MNI |  |  |  | 3700
1090 | T100 Thermal Cycler_BioRad | MNI |  |  |  | 3710
1009 | T2Dx Instrument_T2 Biosystems | MNI |  |  |  | 3720
1086 | T5005-3205_Benchmark Scientific | MNI |  |  |  | 3730
1244 | Talis One instrument_Talis Biomedical | MNI |  |  |  | 3740
1122 | Techne 5Prime/02_Techne | MNI |  |  |  | 3750
1123 | Thermo Fisher QS1_Thermo Fisher | MNI |  |  |  | 3760
1162 | UniCel DxC 600i Synchron Access_BeckmanCoulter | MNI |  |  |  | 3770
1163 | UniCel DxC 660i Synchron Access_BeckmanCoulter | MNI |  |  |  | 3780
1164 | UniCel DxC 680i Synchron Access_BeckmanCoulter | MNI |  |  |  | 3790
1165 | UniCel DxC 860i Synchron Access_BeckmanCoulter | MNI |  |  |  | 3800
1166 | UniCel DxC 880i Synchron Access_BeckmanCoulter | MNI |  |  |  | 3810
1167 | UniCel DxI 600 Access System_BeckmanCoulter | MNI |  |  |  | 3820
1168 | UniCel DxI 800 Access System_BeckmanCoulter | MNI |  |  |  | 3830
1253 | UOL COVID-19 Instrument_Uh Oh Labs | MNI |  |  |  | 3840
1124 | Veriti 96 Fast Thermal Cycler_Thermo Fisher | MNI |  |  |  | 3850
1125 | VIDAS 3_bioMérieux | MNI |  |  |  | 3860
1126 | VIDAS_bioMérieux | MNI |  |  |  | 3870
1006 | VITROS 3600_Ortho-Clinical | MNI |  |  |  | 3880
1005 | VITROS 5600_Ortho-Clinical | MNI |  |  |  | 3890
1008 | VITROS ECi_Ortho-Clinical | MNI |  |  |  | 3900
1007 | VITROS ECiQ_Ortho-Clinical | MNI |  |  |  | 3910
1004 | VITROS XT 7600_Ortho-Clinical | MNI |  |  |  | 3920
| ViraMed Biotech AG_ViraChip Reader | MNI
| LIAISON XS Analyzer_DiaSorin | MNI

## Release Notes
Publication Version ID Last Changed | Change Description | Manufacturer | Model | Vendor Analyte Name | Vendor Specimen Description | Vendor Result Description | Test Performed LOINC Code | Test Performed LOINC Long Name | Test Ordered LOINC Code | Test Ordered LOINC Long Name | Vendor Comment | Vendor Analyte Code | Vendor Reference ID | Testkit Name ID | Testkit Name ID Type | Equipment UID | Equipment UID Type | Component | Property | Time | System | Scale | Method | Publication Version ID | LOINC Version ID | Test ID
20240703 | LIAISON® Lyme Total Antibody Plus can also be run on the LIAISON XS Analyzer. Added a second row to account for additional instrument. | DiaSorin | LIAISON® Lyme Total Antibody Plus | Lyme Total Ab CLIA | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Index = < 0.90 = Negative = Absence of detectable Borrelia burgdorferi antibodies. A negative result does not exclude the possibility of Borrelia burgdorferi infection. If early Lyme disease is suspected, a second sample should be collected and tested two to four weeks later. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Index = ≥ 0.90 and < 1.10 = Equivocal = The imprecision inherent in this method does not allow definitive categorization of samples that read close to the cutoff. Current testing guidelines recommend that all equivocal samples be tested further in a standardized western blot assay. (42425007^Equivocal^SCT) Index = ≥ 1.10 = Positive = Presence of detectable Borrelia burgdorferi antibodies. Current testing guidelines recommend that all positive results be supplemented by further testing in a standardized western blot Assay. (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 101354-9 | Borrelia afzelii+burgdorferi+garinii IgG+IgM Ab [Presence] in Serum by Immunoassay | 101404-2 | Borrelia afzelii+burgdorferi+garinii Ab.IgG and IgM panel - Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme Total Antibody Plus_DiaSorin Inc | MNT | LIAISON XL Analyzer_DiaSorin | MNI | Borrelia afzelii+burgdorferi+garinii Ab.IgG+IgM | PrThr | Pt | Ser | Ord | IA |  | 2.74
20240703 | Updated the result description language for equivocal and positive results per updated package insert. Removed: 'Current testing guidelines recommend that all equivocal samples be tested further in a standardized western blot assay' and 'Current testing guidelines recommend that all positive results be supplemented by further testing in a standardized western blot Assay'. Added: 'All equivocal samples should be tested further following the Standard Two-Tier Testing (STTT) or Modified Two-Tier Testing (MTTT) Methodologies' and 'All positive results should be tested further following the the Standard Two-Tier Testing (STTT) or Modified Two-Tier Testing (MTTT) Methodologies.' | DiaSorin | LIAISON® Lyme Total Antibody Plus | Lyme Total Ab CLIA | Serum (119364003^Serum specimen^SCT) Plasma (119361006^Plasma specimen^SCT) | Index = < 0.90 = Negative = Absence of detectable Borrelia burgdorferi antibodies. A negative result does not exclude the possibility of Borrelia burgdorferi infection. If early Lyme disease is suspected, a second sample should be collected and tested two to four weeks later. (260385009^Negative^SCT - if allowable should map to the recommended term for IA, which is 260415000^Not detected^SCT) Index = ≥ 0.90 and < 1.10 = Equivocal = The imprecision inherent in this method does not allow definitive categorization of samples that read close to the cutoff. Current testing guidelines recommend that all equivocal samples be tested further in a standardized western blot assay. (42425007^Equivocal^SCT) Index = ≥ 1.10 = Positive = Presence of detectable Borrelia burgdorferi antibodies. Current testing guidelines recommend that all positive results be supplemented by further testing in a standardized western blot Assay. (10828004^Positive^SCT - if allowable should map to the recommended term for IA, which is 260373001^Detected^SCT) | 101354-9 | Borrelia afzelii+burgdorferi+garinii IgG+IgM Ab [Presence] in Serum by Immunoassay | 101404-2 | Borrelia afzelii+burgdorferi+garinii Ab.IgG and IgM panel - Serum or Plasma by Immunoassay |  |  |  | LIAISON Lyme Total Antibody Plus_DiaSorin Inc | MNT | LIAISON XL Analyzer_DiaSorin | MNI | Borrelia afzelii+burgdorferi+garinii Ab.IgG+IgM | PrThr | Pt | Ser | Ord | IA |  | 2.74
ADD 1 row for each added test